REAGEN™

Fluoroquinolone ELISA Test Kit

RND99009

### Product Description

REAGEN™Fluoroquinolone ELISA Test Kit enables government agencies, food manufacturers and quality assurance organizations to detect major fluoroquinolones to as low as 0.5 ppb level in various sample types and to satisfy customer concerns about food safety. Fluoroquinolone ELISA Test Kit iprovides a unique extraction method, which is accurate, rapid, cost effective and easy to use. .

The unique features of the kit are:

- The extraction of fluoroquinolones from egg, honey, feed, meat, milk, tissue, serum and    urine can be finished in 10 - 30 minutes with a high recovery rate of 75 - 105%.

- High reproducibility.

- A quick ELISA assay (less than 2 hours regardless of number of  samples..

### Procedure Overview

The method is based on a competitive colorimetric ELISA assay. The drug of interest has been coated in the plate wells. During the analysis, sample is added along with the primary antibody specific for the target drug. If the target is present in the sample, it will compete for the antibody, thereby preventing the antibody from binding to the drug attached to the well. The secondary antibody, tagged with a peroxidase enzyme, targets the primary antibody that is complexed to the drug coated on the plate wells. The resulting color intensity, after addition of substrate, has an inverse relationship with the target concentration in the sample.

### Kit Contents, Storage and Shelf Life

REAGEN™ Fluoroquinolone ELISA Test Kit has the capacity for 96 determinations or testing of 42 samples in duplicate (assuming 12 wells for standards). Return any unused microwells to the foil bag and reseal them with desiccant provided in the original package. Store the kit at 2-8°C*. The shelf life is 12 months when the kit is properly stored.

| Kit Contents | Amount | Storage |
| --- | --- | --- |
| Fluoroquinolone-coated Plate | 1 x 96-well plate (8 wells x 12 strips) | 2-8°C |
| Enrofloxacin Standards: Negative control (white cap tube) 0.1 ng/mL (yellow cap tube) 0.25 ng/mL (orange cap tube) 0.5 ng/mL (pink cap tube) 1.0 ng/mL (purple cap tube) 5.0 ng/mL (blue cap tube) 10 ng/mL (spiking, optional, red cap tube) | 1.0 mL 1.0 mL 1.0 mL 1.0 mL 1.0 mL 1.0 mL 1.0 mL | 2-8°C |
| Fluoroquinolone Antibody #1 | 15 mL | 2-8°C |
| 100X HRP-Conjugated Antibody #2 | 250 µL |  |
| Antibody #2 Diluent ** | 20 mL |  |
| 10X Sample Extraction Buffer ** | 25 mL |  |
| 20X Wash Solution ** | 30 mL |  |
| Stop Buffer ** | 14 mL |  |
| TMB Substrate ** | 10 mL |  |

*If you are not planning to use the kit for over 3 months, store  Fluoroquinolone Antibody #1 and  100X HRP-Conjugated Antibody #2 at -20°C preferably or in a freezer.

Sensitivity (Detection Limit)

| Sample Type | Detection Limit (ng/g or ppb) |
| --- | --- |
| Egg | 1.0 |
| Honey | 2.0 |
| Feed | 1.0 |
| Meat/Liver/Kidney/Fish/Shrimp | 1.0 |
| Milk | 0.5 |
| Serum | 0.5 |
| Urine | 0.5 |

| Fluoroquinolines | Cross-Reactivity (%) |
| --- | --- |
| Enrofloxacin | 100 |
| Ciprofloxacin | 102 |
| Danofloxacin | 90 |
| Sarafloxacin | 96 |
| Norfloxacin | 67 |
| Enoxacin | 75 |
| Pipemidic acid | 31 |
| Ofloxacin | 72 |
| Lomefloxacin | 86 |
| Difloxacin | 90 |
| Pefloxacin | 69 |
| Benofloxacin | 42 |
| Flumequin | 62 |
| Oxolin acid | 91 |

Specificity (Cross-Reactivity)

### Required Materials Not Provided With the Kit

- Microtiter plate reader (450 nm)

- Incubator

- Tissue Mixer (e.g. Omni TissueMaster Homogenizer)

- Vortex mixer (e.g. Gneie Vortex mixer from VWR)

- 10, 20, 100 and 1000 L pipettes

- Multi-channel pipette: 50-300 L (Optional)

### Warnings and Precautions

- Do not use the kit past the expiration date.

- Do not intermix reagents from different kits or lots except for components with the same part No’s within their expiration dates.

- Try to maintain a laboratory temperature of 20°–25°C (68°–77°F). Avoid running assays under or near air vents, as this may cause excessive cooling, heating and/or evaporation. Also, do not run assays in direct sunlight, as this may cause excessive heat and evaporation. Cold bench tops should be avoided by placing several layers of paper towel or some other insulation material under the assay plates during incubation.

- Make sure you are using only distilled or deionized water since water quality is very important.

- When pipetting reagents into an empty microtiter plate, place the pipette tips

in the lower corner of the well, making contact with the plastic.

- Incubations of assay plates should be timed as precisely as possible. Be consistent when adding standards to the assay plate. Add your standards first and then your samples.

- Add standards to plate only in the order from low concentration to high concentration as this will minimize the risk of compromising the standard curve.

- Always refrigerate plates in sealed bags with a desiccant to maintain stability. Prevent condensation from forming on plates by allowing them equilibrate to room temperature (20 – 25C / 68 – 77F) while in the packaging.

SAMPLE PREPARATION

Be sure samples are properly stored. In general, samples should be refrigerated at 2-4°C for no more than 1-2 days. Freeze samples to a minimum of -20°C if they need to be stored for a longer period. Frozen samples can be thawed at room temps (20 – 25C / 68 – 77F) or in a refrigerator before use.

Egg

    - Weigh out 1.0 g of egg sample and add 4 mL of 70% methanol.

    - Vortex for 2 minutes at maximum speed.

    - Centrifuge for 5 minutes at 4,000 x g at room temperature (20 – 25°C / 68 – 77°F).

    - Transfer 0.5 mL of the supernatant to a tube, add 0.5 mL of 1X Sample Extraction Buffer, and    mix well.

    - Use 50 µL of the sample per well for the assay.

Note: Dilution factor: 10.

Feed

1.  To a reasonable amount of sample, add 5 times the amount of 70%  methanol (e.g. 1 g of sample + 5 mL of 70% methanol); homogenize the sample with a suitable mixer.

2.  Take out 1.5 mL of the homogenized sample and centrifuge for 5 minutes at 4,000 x g at room temperature (20 – 25°C / 68 – 77°F).

3.  Transfer 0.5 mL of the supernatant to a tube, add 0.5 mL of 1X Sample Extraction Buffer, and mix well.

4.  Use 50 µL of the sample per well for the assay.

Note: Dilution factor: 10.

Meat/Liver/Kidney/Fish/Shrimp

    - emove fat from the sample. Homogenize the sample with a suitable mixer.

2.    Weigh out 1.0 g of the homogenized sample and add 4 mL of 70% methanol.

3 .   Vortex for 10 minutes at maximum speed.

4.  Centrifuge for 5 minutes at 4,000 x g at room temperature (20 – 25°C / 68 – 77°F).

5.    Transfer 0.5 mL of the supernatant to a tube, add 0.5 mL of 1X Sample Extraction Buffer, and mix well.

6.  Use 50 µL for the assay.

Note: Dilution factor: 10.

Milk

1.  For fat-free milk, take out 200 µL of the milk sample, add 800 µL of 35% Methanol/Sample

Extraction Buffer, and mix well. Take 50 µL of the diluted sample per well for the assay.

2.  For regular milk with fat, centrifuge 1 mL of the milk sample at 4,000 x g for 5 minutes, discard the upper fat layer. Take out 200 µL  of  the milk  sample, add 800 µL  of  35% Methanol/Sample Extraction Buffer, and mix well. Take 50 µL of the diluted sample per well for the assay.

Note: Dilution factor: 5.

Serum

- Centrifuge 0.5 mL of urine at 4,000 x g for 5 minutes.

2.   Take  out 0.2 mL  of  the supernatant,  add 0.8 mL  of  35% Methanol/Sample  Extraction Buffer.

3.   Use 50 µL per well for the assay.

Note: Dilution factor: 5.

Urine

1.  Centrifuge 0.5 mL of urine at 4,000 x g for 5 minutes.

2.  Take  out 0.2 mL  of  the supernatant,  add 0.8 mL  of  35%  Methanol/Sample  Extraction

3.  Use 50 µL per well for the assay.

Note: Dilution factor: 5

Honey

1.  Take out 0.1 g of honey sample, add 1.9 mL of 35%  Methanol/Sample Extraction Buffer and mix well.

2.  Use 50 µL per well for the assay.

Note: Dilution factor: 20

FLUOROQUINOLONE  ELISA TEST KIT PROTOCOL

### Reagent Preparation

IMPORTANT: All reagents should be brought up to room temperature before use (1 – 2 hours at 20 – 25C / 68 – 77F); Make sure you read “Warnings and Precautions” section on page 3. Solutions should be prepared just prior to ELISA test.  All reagents should be mixed by gently inverting or swirling prior to use. Prepare volumes that are needed for the number of wells being run. Do not return the reagents to the original stock tubes/bottles. Using disposable reservoirs when handling reagents can minimize the risk of contamination and is recommended.

1.  Preparation of 1X HRP-Conjugated Antibody #2

Mix 1 volume of 100X HRP-Conjugated Antibody #2 with 99 volumes of Antibody #2 Diluent.

2.  Preparation of 1X Wash Solution

Mix 1 volume of the 20X Wash Solution with 19 volumes of distilled water.

3.	Preparation of 1X Sample Extraction Buffer

Mix 1 volume of 10X Sample Extraction Buffer with 9 volumes of distilled water.

4.		Preparation of 35% Methanol/Sample  Extraction  Buffer (For tests on Milk, Serum and Urine samples only)

Mix 6.5  volumes of 1X Sample Extraction Buffer with 3.5 volumes of 100%  methanol.

### ELISA Testing Protocol

Label the individual strips that will be used and aliquot reagents as the following example:

| Component | Volume per Reaction | 24 Reactions |
| --- | --- | --- |
| Fluoroquinolone Antibody #1 | 100 µL | 2.4 mL |
| 1X HRP-Conjugated Antibody #2 | 150 µL | 3.6 mL |
| 1X Wash Solution | 2.0 mL | 48 mL |
| Stop Buffer | 100 µL | 2.4 mL |
| TMBSubstrate | 100 µL | 2.4 mL |

1. Add 50 µL of each Enrofloxacin Standards in duplicate into different wells (Add standards to plate only in the order from low concentration to high concentration).

2. Add 50 µL of each sample in duplicate into different sample wells.

3. Add 100 µL of Antibody #1 and mix well by gently rocking the plate manually for 1 minute.

4. Incubate the plate for 30 minutes at room temperature (20 – 25°C / 68 – 77°F).

5. Wash the plate 3 times with 250 µL of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels (Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

6. Add 150 µL of 1X HRP-Conjugated Antibody #2. Incubate the plate for 30 minutes at room temperature (20 – 25°C / 68 – 77°F) (Avoid direct sunlight and cold bench tops during the incubation. Covering the microtiter plate while incubating is recommended).

7. Wash the plate 3 times with 250 µL of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels (Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

8. Add 100 µL of TMB substrate. Time the reaction immediately after adding the substrate.Mix the solution by gently rocking the plate manually for 1 minute while incubating

9. After incubating for 15 minutes at room temperature (20 – 25°C / 68 – 77°F), add 100 µLof Stop Buffer to stop the enzyme reaction.

10. Read the plate as soon as possible following the addition of Stop  Buffer on a plate reader with 450 nm wavelength (Before reading, use a lint-free wipe on the bottom of the plate to ensure no moisture or fingerprints interfere with the readings).

### Fluoroquinolone Concentration Calculations

A standard curve can be constructed by plotting the mean relative absorbance (%) obtained from each reference standard against its concentration in ng/mL on a logarithmic curve.

Relative absorbance (%) =

The following figure is a typical fluoroquinolones standard curve

REAGEN LLC

Add.: 102CommerceDr.Unit8,Moorestown,NJ08057

Tel.:856-727-0250         FAX.:856-727-0251

E-mail:reagenkits@gmail.com  reagenllc@gmail.com

Web.: www.reagen.us