REAGEN™

Apramycin ELISA Test Kit

RND99093

### Product Description

REAGEN™ Apramycin ELISA Test Kit is a competitive enzyme immunoassay for the quantitative analysis of apramycin in chicken, fish, meat, liver, and kidney.

The unique features of the kit are:

- High recovery (>80%), rapid (10 - 40 minutes), and cost-effective extraction methods.

- High sensitivity (1.5 ng/g or ppb)

- A quick ELISA assay (less than 2 hours regardless of number of samples).

- High reproducibility.

### Procedure Overview

The method is based on a competitive colorimetric ELISA assay. The drug of interest has been coated in the plate wells. During the analysis, sample is added along with the enzyme-conjugate and the primary antibody specific for the target drug. If the target is present in the sample, it will compete for the antibody, thereby preventing the antibody from binding to the drug attached to the well. The enzyme-conjugate, will bind with the primary antibody that is complexed to the drug coated on the plate wells. The resulting color intensity, after addition of substrate, has an inverse relationship with the target concentration in the sample

### Kit Contents, Storage and Shelf Life

| Kit Contents | Amount | Storage |
| --- | --- | --- |
| Apramycin Coated Plate | 1 x 96-well Plate (8 wells x 12 strips) | 2-8C |
| Apramycin Standards: Negative control (white cap tube) 0.25 ng/mL (yellow cap tube) 1.0 ng/mL (orange cap tube) 2.5 ng/mL (pink cap tube) 5.0 ng/mL(purple cap tube) 15 ng/mL (blue cap tube) Spiking 1000 ng/mL(red cap tube) | 0.8ml 0.8ml 0.8ml 0.8ml 0.8ml 0.8ml 0.8ml | 2-8C * |
| Apramycin Antibody #1 | 12 mL | 2-8C * |
| TMB Substrate ** | 10 mL |  |
| 100X HRP-Conjugated Antibody #2 | 250 uL |  |
| Antibody #2 Diluent ** | 20 mL |  |
| Stop Buffer ** | 20 mL |  |
| 20X Wash Solution ** | 30 mL |  |
| 10X Apramycin Extraction Buffer* | 28 mL |  |

* If you are not planning to use the kit for over 1 month, store Apramycin Standard Stock,

Apramycin Antibody #1 and 100X HRP-Conjugated Antibody #2 at -20℃ or in a freezer.

### Sensitivity

| Antibiotic | Chicken/Fish/Meat/Liver/Kidney(ppb) |
| --- | --- |
| Apramycin | 10 |

### Specificity

| Analytes | Cross-Reactivity (%) |
| --- | --- |
| Apramycin | 100 |

### Required Materials Not Provided With the Kit

- Microtiter plate reader (450 nm)

- Incubator

- Tissue Mixer (e.g. Omni TissueMaster Homogenizer)

- Vortex mixer (e.g. Gneie Vortex mixer from VWR)

- 10, 20, 100 and 1000 uL pipettes

- Multi-channel pipette: 50-300 uL (Optional)

- n-Hexane

- Distilled deionized water

- Disposable plastic reservoirs

- Paper towels

### Warnings and Precautions

- The standards contain Apramycin. Handle with particular care.

- Do not use the kit past the expiration date.

- Do not intermix reagents from different kits or lots except for components with the same part No’s within their expiration dates. ANTIBODIES AND PLATES ARE KIT-AND LOT-SPECIFIC. Make sure that the antibody #2 and diluent are mixed in correct volumes.

- Try to maintain a laboratory temperature of 20°–25°C (68°–77°F). Avoid running assays under or near air vents, as this may cause excessive cooling, heating and/or evaporation. Also, do not run assays in direct sunlight, as this may cause excessive heat and evaporation. Cold bench tops should be avoided by placing several layers of paper towel or some other insulation material under the assay plates during incubation.

- Make sure you are using only distilled or deionized water since water quality is very important.

- When pipetting samples or reagents into an empty microtiter plate, place the pipette tips in the lower corner of the well, making contact with the plastic.

- Incubations of assay plates should be timed as precisely as possible. Be consistent when adding standards to the assay plate. Add your standards first and then your samples.

- Add standards to plate only in the order from low concentration to high concentration as this will minimize the risk of compromising the standard curve.

- Always refrigerate plates in sealed bags with a desiccant to maintain stability. Prevent condensation from forming on plates by allowing them equilibrate to room temperature (20 – 25C / 68 – 77F) while in the packaging.

SAMPLE PREPARATION

Be sure samples are properly stored. In general, samples should be refrigerated at 2-4°C for no more than 1-2 days. Freeze samples to a minimum of -20°C if they need to be stored for a longer period. Frozen samples can be thawed at room temps (20 – 25C / 68 – 77F) or in a refrigerator before use.

- Preparation of 1X Apramycin Extraction Buffer:

Mix 1 volume of 10X Apramycin Buffer with 9 volumes of distilled water.

- Preparation of 35% Ethanol/1X Sample Extraction Buffer

Mix 13 volume of 1X Sample Extraction Buffer with 7 volumes of pure ethanol. Vortex for 2

minutes prior to use.

### Chicken/Fish/ Meat/Liver/Kidney

    - To 1.0 g of the homogenized sample, add 4mL of 35% Ethanol/1X Sample Extraction Buffer. Vortex the sample for 3minutesmanually or vortex the sample with multi-vortexer for 10minutes.

    - Centrifuge the sample for 10 minutes at 2,000 x g.

    - Carefully transfer 100 uL of the supernatant to a new tube, add 700 uL of 35% Ethanol/1X Sample Extraction Buffer. Vortex for 60 seconds.

    - Use 50 uLper well for the assay.

Note:Dilution factor: 40.

Apramycin  ELISA TEST KIT PROTOCOL

### Reagent Preparation

IMPORTANT: All reagents should be brought up to room temperature before use (1 – 2 hours at 20 – 25C / 68 – 77F); Make sure you read “Warnings and Precautions” section on page 3. Solutions should be prepared just prior to ELISA test.  All reagents should be mixed by gently inverting or swirling prior to use. Prepare volumes that are needed for the number of wells being run. Do not return the reagents to the original stock tubes/bottles. Using disposable reservoirs when handling reagents can minimize the risk of contamination and is recommended.

- Preparation of 1X HRP-Conjugated Antibody #2

Mix 1 volume of 100X HRP-Conjugated Antibody #2 with 199 volumes of Antibody #2 Diluent.

- Preparation of 1X Wash Solution

Mix 1 volume of the 20X Wash Solution with 19 volumes of distilled water.

### ELISA Testing Protocol

Label the individual strips that will be used and aliquot reagents as the following example:

| Component | Volume per Reaction | 24 Reactions |
| --- | --- | --- |
| 1X Apramycin Antibody #1 | 100 uL | 2.4 mL |
| 1X HRP-Conjugated Antibody #2 | 150 uL | 3.6 mL |
| 1X Wash Solution | 2.0 mL | 48 mL |
| Stop Buffer | 100 uL | 2.4 mL |
| TMB Substrate | 100 uL | 2.4 mL |

    - Add 50 uL of each Apramycin Standards in duplicate into different wells ( Add standards to plate only in the order from low concentration to high concentration).

    - Add 50 uL of each sample in duplicate into different sample wells.

    - Add 100 uL of 1X Apramycin  Antibody #1 and mix well by gently rocking the plate manually for 1 minute.

    - Incubate the plate for 30 minutes at room temperature (20 – 25C / 68 – 77F). 

    - Wash the plate 3 times with 250 uL of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels ( Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

    - Add 150 uL of 1X Antibody #2 solution. Incubate the plate for 30 minutes at room temperature (20 – 25C / 68 – 77F) (  Avoid direct sunlight and cold bench tops during the incubation. Covering the microtiter plate while incubating is recommended).

    - Wash the plate 3 times with 250 uL 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels ( Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

    - Add 100 uL of TMB substrate. Time the reaction immediately after adding the substrate. Mix the solution by gently rocking the plate manually for 1 minute while incubating (  Do not put any substrate back to the original container to avoid any potential contamination. Any substrate solution exhibiting coloration is indicative of deterioration and should be discarded. Covering the microtiter plate while incubating is recommended).

    - After incubating for 10-15 minutes at room temperature (20 – 25C / 68 – 77F), add 100 uL of Stop Buffer to stop the enzyme reaction.

    - Read the plate as soon as possible following the addition of Stop Buffer on a plate reader with 450 nm wavelength ( Before reading, use a lint-free wipe on the bottom of the plate to ensure no moisture or fingerprints interfere with the readings).

### Apramycin Concentration Calculations

A standard curve can be constructed by plotting the mean relative absorbance (%) obtained from each reference standard against its concentration in ng/mL on a logarithmic curve.

Relative absorbance (%) =

Use the mean relative absorbance values for each sample to determine the corresponding concentration of the tested drug in ng/mL from the standard curve.

The following figure is a typical chloramphenicol standard curve.

REAGEN LLC

Add.: 102CommerceDr.Unit8,Moorestown,NJ08057

Tel.:856-727-0250         FAX.:856-727-0251

E-mail:reagenkits@gmail.com  reagenllc@gmail.com

Web.: www.reagen.us