REAGEN™

Amantadine ELISA Test Kit

RND99071

### Product Description

REAGEN™ Amantadine ELISA Test Kit is a competitive enzyme immunoassay for the quantitative analysis of amantadine in eggs

The unique features of the kit are:

- High recovery(>75%), rapid (less than 60 minutes), and cost-effective extraction methods.

- High sensitivity (0.25 ng/g or ppb)

- High reproducibility.

- A quick ELISA assay (less than 2 hours regardless of number of samples).

### Procedure Overview

REAGEN™ Amantadine ELISA Test Kit is based on a competitive colorimetric ELISA assay. The plate wells are coated with anti-amantadine antibody. During analysis, sample is added along with the Amantadine-horseradish peroxidase (Amantadine-HRP) conjugate. If Amantadine residue is present in the sample, it will compete for the Amantadine antibody, thereby preventing the Amantadine-HRP from binding to the antibody attached to the well. The resulting color intensity, after addition of the HRP substrate (TMB), has an inverse relationship with the concentration of Amantadine residue in the sample.

### Kit Contents, Storage and Shelf Life

REAGEN™ Amantadine ELISA Test Kit has the capacity for 96 determinations or testing of 42 samples in duplicate (assuming 12 wells for standards). Return any unused microwells to the foil bag and reseal them with the desiccant provided in the original package. Store the kit at 2-8C . The shelf life is 12 months when the kit is properly stored.

| Kit Contents | Amount | Storage |
| --- | --- | --- |
| Amantadine Ab-Coated Plate | 1 x 96-well plate (8 wells x 12 strips) | 2-8℃ |
| Amantadine Standards Negative control (white CAP tube) 0.5 ng/mL (yellow CAP tube) 1.5 ng/mL (orange CAP tube) 4.5ng/mL (pink CAP tube) 13.5ng/mL (purple CAP tube) 27 ng/mL (blue CAP tube) 100 ng/mL(spiking, optional, red cap tube) | 1mL 1mL 1mL 1mL 1mL 1mL 1mL | 2-8℃ |
| Amantadine-HRP Conjugate | 6 mL | 2-8℃ |
| 20× Wash Solution | 28mL |  |
| Stop Buffer | 14mL |  |
| TMB Substrate | 12mL |  |
| 4X Egg Sample Suspension Buffer | 10 mL |  |
| 10XEgg Powder Sample Suspension Buffer | 6 mL |  |
| Amantadine Extraction Additive I | 8 mL |  |
| Amantadine Extraction Additive II | 20 mL |  |
| Amantadine Balance Buffer I | 20 mL |  |
| Amantadine Balance Buffer II | 62g |  |

If you are not planning to use the kit for over 1 month , HRP-Conjugated at 20℃ or in a freezer is recommended.

### Sensitivity

| Sample Type | Detection Limit (ng/g or ppb) |
| --- | --- |
| egg | 0.5 |
| Egg powder | 0.5 |

### Specificity

| Analytes | Cross-Reactivity (%) |
| --- | --- |
| Amantadine | 100 |

### Required Materials Not Provided With the Kit

- Microtiter plate reader (450 nm)

- Constant temperature incubator

- Vortex mixer, (e.g. Gneie Vortex mixer from VWR)

- Tissue Mixer (e.g. Omni Tissue Master Homogenizer)

- Pressure Blowing Concentrator

- 10, 20, 100 and 1000 uL pipettes

- Multi-channel pipette: 50-300 uL (Optional)

### Warnings and Precautions

- The standards contain Amantadine. Handle with particular care.

- Do not use the kit past the expiration date.

- Do not intermix reagents from different kits or lots except for components with the same part No’s within their expiration dates. PLATES ARE KIT-AND LOT-SPECIFIC.

- Try to maintain a laboratory temperature of 20–25°C (68°–77°F). Avoid running assays under or near air vents, as this may cause excessive cooling, heating and/or evaporation. Also, do not run assays in direct sunlight, as this may cause excessive heat and evaporation. Cold bench tops should be avoided by placing several layers of paper towel or some other insulation material under the assay plates during incubation.

- Make sure you are using only distilled or deionized water since water quality is very important.

- When pipetting samples or reagents into an empty microtiter plate, place the pipette tips in the lower corner of the well, making contact with the plastic.

- Incubations of assay plates should be timed as precisely as possible. Be consistent when adding standards to the assay plate. Add your standards first and then your samples.

- Add standards to plate only in the order from low concentration to high concentration as this will minimize the risk of compromising the standard curve.

- Always refrigerate plates in sealed bags with a desiccant to maintain stability. Prevent condensation from forming on plates by allowing them equilibrate to room temperature (20 – 25℃ / 68 – 77℉) while in the packaging.

SAMPLE PREPARATION

Be sure samples are properly stored. In general, samples should be refrigerated at 2-4℃ for no more than 1-2 days. Freeze samples to a minimum of -20℃ if they need to be stored for a longer period. Frozen samples can be thawed at room temps (20 – 25℃ / 68 – 77℉) or in a refrigerator before use.

- Preparation of 1X Egg Sample Suspension Buffer

Mix 1 volume of 4X Egg Sample Suspension Buffer with 3 volumes of distilled water.

- Preparation of 1X Egg Powder Sample Suspension Buffer

Mix 1 volume of 10X Egg Powder Sample Suspension Buffer with 9 volumes of distilled water.

- Preparation of 1X Amantadine Balance Buffer

Add 20 mL of the Amantadine Balance Buffer I and 62 g of Amantadine Balance Buffer II (Powder) to a 250-mL plastic bottle. Add 180 mL of distilled water to the bottle, mix well to obtain a homogeneous solution. (To speed up the powder dissolving, the solution can be heated at 60°C during this process. The prepared solution can be store at room temperature for one year).

- Preparation of Acetonitrile Containing 1% of Amantadine Extraction Additive I

Mix 99 volumes of acetonitrile with 1 volume of Amantadine Extraction Additive I.

Caution: Amantadine Extraction Additive I contains a strong acidic solution! Always wear gloves and personal protective equipment when preparing and handling this buffer during sample preparation!  Prepare this buffer in a fume hood!

egg

- Homogenize a reasonable amount of sample with a suitable mixer.

- To 2 g of a homogenized sample in a tube, add 6 mL of acetonitrile containing 1% of Amantadine Extraction Additive I and 3 mL of hexane. Manually vortex for 3 minutes at maximum speed, or vortex for 10 minutes using a multi-tube vortexer.

- Centrifuge for 5 minutes at 4,000 x g at room temperature (20 – 25 °C / 68 – 77 °F).

- Remove the top hexane layers completely. Transfer 3 mL supernatant to another 15 mL conical tube, add 2 mL of 1X Amantadine Balance Buffer to each tube first, and then shake all tubes together for 15 seconds.  Add 200 uL of Amantadine Extraction Additive II to each tube. Vortex for 2 minutes at maximum speed.

- Centrifuge for 5 minutes at 4,000 x g at room temperature (20 – 25 °C / 68 – 77 °F).

- Slowly transfer 1.6 mL of the acetonitrile supernatant (corresponding to 0.4 g of the original sample) into a new vial and use a rotary evaporator to dry the sample in a 60 – 70 °C water bath under reduced pressure. Alternatively, the sample can be dried by blowing nitrogen gas over it in a 60 – 70 °C water bath.

- To the dried residue, add 0.4 mL of 1X Egg Sample Suspension Buffer. Vortex at maximum speed for 2 minutes.

- Use 100 uL of the sample per well for the assay.

Note: Dilution factor: 1

Egg Powder

- To 2 g of a homogenized sample in a tube, add 2 mL distilled water, 6 mL of acetonitrile containing 1% of Amantadine Extraction Additive I, and 3 mL of hexane. Manually vortex for 3 minutes at maximum speed, or vortex for 10 minutes using a multi-tube vortexer.

- Centrifuge for 5 minutes at 4,000 x g at room temperature (20 – 25 °C / 68 – 77 °F).

- Remove the top hexane layers completely. Transfer 3.5 mL supernatant to another 15 mL conical tube, add 2 mL of 1X Amantadine Balance Buffer to each tube first, and then shake all tubes together for 15 seconds.  Add 200 uL of Amantadine Extraction Additive II to each tube. Vortex for 2 minutes at maximum speed.

- Centrifuge for 5 minutes at 4,000 x g at room temperature (20 – 25 °C / 68 – 77 °F).

- Slowly transfer 1.6 mL of the acetonitrile supernatant (corresponding to 0.4 g of the original sample) into a new vial and use a rotary evaporator to dry the sample in a 60 – 70 °C water bath under reduced pressure. Alternatively, the sample can be dried by blowing nitrogen gas over it in a 60 – 70 °C water bath.

- To the dried residue, add 0.4 mL of 1X Egg Powder Sample Suspension Buffer. Vortex at maximum speed for 2 minutes.

- Use 100 uL of the sample per well for the assay.

Note: Dilution factor = 1.

AMANTADINE ELISA TEST KIT PROTOCOL

### Reagent Preparation

IMPORTANT: All reagents should be brought up to room temperature before use (1 – 2 hours at 20 – 25°C / 68 – 77℉); Make sure you read “Warnings and Precautions” section. Solutions should be prepared just prior to ELISA test. All reagents should be mixed by gently  inverting or swirling prior to use. Prepare volumes that are needed for the number of wells being run. Do not return the reagents to the original stock tubes/bottles. Using disposable reservoirs when handling reagents can minimize the risk of contamination and is recommended.

- Preparation of 1X Wash Solution

Mix 1 volume of the 20X Wash Solution with 19 volumes of distilled water.

### ELISA Testing Protocol

Label the individual strips that will be used and aliquot reagents as the following example:

| Component | Volume per Reaction | 24 Reactions |
| --- | --- | --- |
| HRP-Conjugated | 50 uL | 1.2 mL |
| 1X Wash Solution | 2.0mL | 48 mL |
| Stop Buffer | 100 uL | 2.4 mL |
| TMB Substrate | 100 uL | 2.4 mL |

    - Add 100 uL of each Amantadine Standards and sample in duplicate into different sample wells.

    - Add 50 uL of HRP-Conjugated and mix well by gently rocking the plate manually for 1 minute.

    - Incubate the plate for 30 minutes at room temperature (20 – 25℃ / 68 – 77℉).

    - Wash the plate 3 times with 250 uL of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels ( Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

    - Add 100 uL of TMB Substrate to each well. Time the reaction immediately after adding the substrate. Mix the solution by gently rocking the plate manually for 1 minute while incubating ( Do not put any substrate back to the original container to avoid any potential contamination. Covering the microtiter plate while incubating is recommended).

    - After incubating for 15 minutes at room temperature (20 – 25°C / 68 – 77℉), add 100 uL of Stop Buffer to stop the enzyme reaction.

    - Read the plate as soon as possible following the addition of Stop Buffer on a plate reader with 450 nm wavelength ( Before reading, use a lint-free wipe on the bottom of the plate to ensure no moisture or fingerprints interfere with the readings).

### Amantadine Concentration Calculations

A standard curve can be constructed by plotting the mean relative absorbance (%) obtained from each reference standard against its concentration in ng/mL on a logarithmic curve.

Relative absorbance (%) =

The following figure is a typical Amantadine standard curve.

REAGEN LLC

Add.: 102CommerceDr.Unit8,Moorestown,NJ08057

Tel.:856-727-0250         FAX.:856-727-0251

E-mail:reagenkits@gmail.com  reagenllc@gmail.com

Web.:www.reagen.us