REAGEN™

Dimetridazole ELISA Test Kit

RND99062

### Product Description

REAGEN™ Dimetridazole ELISA Test Kit provides a competitive enzyme immunoassay for the quantitative analysis of dimetridazole in meat/liver/kidney, Fish ,shrimp, honey ,egg , feed and urine.

The unique features of the kit are:

- High recovery (75%-115%), and cost-effective extraction methods for various samples.

- A quick ELISA assay (less than 1.5 hours regardless of number of samples).

- High reproducibility.

### Procedure Overview

The method is based on a competitive colorimetric ELISA assay. The drug of interest has been coated in plate wells. During the analysis, sample is added along with the primary antibody specific for target drug. If the target is present in the sample, it will compete for the antibody, thereby preventing the antibody from binding to the drug attached to the well. The second antibody, tagged with a peroxidase enzyme, targets the primary antibody that is complexed to the drug coated on the wells. The resulting color intensity, after addition of substrate, has an inverse relationship with the target concentration in the sample.

### Kit Contents, Storage and Shelf Life

REAGEN™ Dimetridazole ELISA Test Kit has the capacity for 96 determinations or testing of 42 samples in duplicate (assuming 12 wells for standards). Return any unused microwells to the foil bag and reseal them with the desiccant provided in the original package. Store the kit at 2-8C*. The shelf life is 12 months when the kit is properly stored.

| Kit Contents | Amount | Storage |
| --- | --- | --- |
| Dimetridazole -coated Microtiter Plate | 1 x 96-well plate (8 wells x 12 strips) | 2-8C |
| Dimetridazole Standards: Negative control 0.5ng/mL 1.5 ng/mL 3 ng/mL 6 ng/mL 12 ng/mL 100ng/mL (spiking, optional) | 0.8 mL 0.8 mL 0.8 mL 0.8 mL 0.8 mL 0.8 mL 0.8 mL | 2-8C |
| Antibody#1 | 12mL | 2-8C * |
| 100x HRP Antibody#2 | 800 L | 2-8C * |
| HRP Antibody#2 Diluent | 20 mL | 2-8C |
| TMB Substrate | 12mL | 2-8C |
| 20X Wash Solution ** | 28 mL | 2-8C |
| Stop Buffer ** | 14 mL | 2-8C |
| Concentrate of Sample Extraction Buffer | 30 g | 2-8C |
| Sample Dilution Buffer | 20 mL x 2 | 2-8C |

* 	If you are not planning to use the kit for over 1 month, store 100 x HRP Antibody#2 at -20C or in a freezer.

### Sensitivity

| Sample Type | Detection Limit (ng/g or ppb) |
| --- | --- |
| Meat/Liver/Kidney | 0.5 |
| /Fish/Shrimp/Honey | 0.5 |
| Egg | 0.5 |
| Feed | 1.0 |
| Urine | 2.5 |

### Specificity

| Analytes | Cross-Reactivity (%) |
| --- | --- |
| Dimetridazole | 100 |
| Metronidazole | 100 |
| Hydroxydimetridazole | 32 |
| Ronidazole | 14 |
| Ipronidazole | 11 |
| Nicarbazin | 7 |
| Halofuginone | 3 |
| Diclazuril | 4 |
| Robenidine | 2 |

### Required Materials Not Provided With the Kit

- Microtiter plate reader (450 nm)

- Rotary evaporator or nitrogen gas

- oscillator

- Vortex mixer (e.g. Gneie Vortex mixer from VWR)

- 20, 100 and 1000 L pipettes

- Multi-channel pipette: 50-300 L (Optional)

- Acetonitrile

- n-hexane

- deionized water

### Warnings and Precautions

- The standards contain Dimetridazole . Handle with particular care.

- Do not use the kit past the expiration date.

- Do not intermix reagents from different kits or lots except for components with the same part No’s within their expiration dates. HRP AND PLATES ARE KIT- AND LOT-SPECIFIC. Make sure that the antibody #2 and diluent are mixed in correct volumes.

- Try to maintain a laboratory temperature of 20°–25°C (68°–77°F). Avoid running assays under or near air vents, as this may cause excessive cooling, heating and/or evaporation. Also, do not run assays in direct sunlight, as this may cause excessive heat and evaporation. Cold bench tops should be avoided by placing several layers of paper towel or some other insulation material under the assay plates during incubation.

- Make sure you are using only distilled or deionized water since water quality is very important.

- When pipetting samples or reagents into an empty microtiter plate, place the pipette tips in the lower corner of the well, making contact with the plastic.

- Incubations of assay plates should be timed as precisely as possible. Be consistent when adding standards to the assay plate. Add your standards first and then your samples.

- Add standards to plate only in the order from low concentration to high concentration as this will minimize the risk of compromising the standard curve.

- Always refrigerate plates in sealed bags with a desiccant to maintain stability. Prevent condensation from forming on plates by allowing them equilibrate to room temperature (20 – 25C / 68 – 77F) while in the packaging.

SAMPLE PREPARATION

Be sure samples are properly stored. In general, samples should be refrigerated at 2-4°C for no more than 1-2 days. Freeze samples to a minimum of -20°C if they need to be stored for a longer period. Frozen samples can be thawed at room temps (20 – 25C / 68 – 77F) or in a refrigerator before use.

- Preparation of 1X Sample Extraction Buffer

Take all of the powder from the Concentrate of Sample Extraction Buffer bag to a 125-mL bottle, add 90 mL of distilled water, vortex 2 minutes manually, leave the solution at room temperature for 20 minutes. It is fine if small amount of salt residue is observed at the bottom of the bottle.

### Meat/Liver/Kidney

- Remove fat from the sample. Homogenize the sample with a suitable mixer.

- Weigh out 2.0 g of the homogenized sample and add 1 mL of 1X Sample Extraction Buffer into a centrifugal vial.

- Add 4 mL of Acetonitrile, vortex for 3 minutes at maximum speed manually or 10 minutes using a multi-tube vortexer.

- Centrifuge for 5 minutes at 4,000 x g at room temperature (20 – 25C / 68 – 77F).

- Transfer 1.0 mL of the supernatant to another tube (corresponding to 0.5 g of the original sample) into a new vial and use a rotary evaporator to dry the sample in a 60-70C water bath under reduced pressure. Alternatively, the sample can be dried by blowing nitrogen gas in a 60-70C water bath.

- Dissolve the dried residue in 0.5 mL of Sample Dilution Buffer and 1 mL of hexane by vortexing at maximum speed for 1 minute.

- Centrifuge for 5 minutes at 4,000 x g at room temperature (20 – 25C / 68 – 77F).

- Remove top hexane layer completely, use 50 L of the bottom layer sample per well for the assay.

Note: Dilution factor: 1.

### Shrimp/ Fish/ Honey

- Remove fat from the sample. Homogenize the sample with a suitable mixer.

- Weigh out 2.0 g of the homogenized sample and add 1 mL of 1X Sample Extraction  Buffer into a centrifugal vial.

- Add 4 mL of Acetonitrile, vortex for 3 minutes at maximum speed manually or 10 minutes using a multi-tube vortexer.

- Centrifuge for 5 minutes at 4,000 x g at room temperature (20 – 25C / 68 – 77F).

- Transfer 1.0 mL of the supernatant to another tube (corresponding to 0.5 g of the original sample) into a new vial and use a rotary evaporator to dry the sample in a 60-70C water bath under reduced pressure. Alternatively, the sample can be dried by blowing nitrogen gas in a 60-70C water bath.

- Dissolve the dried residue in 0.5 mL of Sample Dilution Buffer and 1 mL of hexane by vortexing at maximum speed for 1 minute.

- Centrifuge for 5 minutes at 4,000 x g at room temperature (20 – 25C / 68 – 77F).

- Remove top hexane layer completely, use 50 L of the bottom layer sample per well for the assay.

Note: Dilution factor: 1.

### Egg

- Remove the eggshell from the sample. Homogenize the sample with a suitable mixer.

- Weigh out 2.0 g of the homogenized sample and add 1 mL of 1X Sample Extraction  Buffer into a centrifugal vial.

- Add 4 mL of Acetonitrile, vortex for 3 minutes at maximum speed manually or 10 minutes using a multi-tube vortexer.

- Centrifuge for 5 minutes at 4,000 x g at room temperature (20 – 25C / 68 – 77F).

- Transfer 1.0 mL of the supernatant to another tube (corresponding to 0.5 g of the original sample) into a new vial and use a rotary evaporator to dry the sample in a 60-70C water bath under reduced pressure. Alternatively, the sample can be dried by blowing nitrogen gas in a 60-70C water bath.

- Dissolve the dried residue in 0.5 mL of Sample Dilution Buffer and 1 mL of hexane by vortexing at maximum speed for 1 minute.

- Centrifuge for 5 minutes at 4,000 x g at room temperature (20 – 25C / 68 – 77F).

- Remove top hexane layer completely, use 50 L of the bottom layer sample per well for the assay.

Note: Dilution factor: 1.

·Feed

- Weigh out 2.0 g of the homogenized sample and add 1 mL of 1X Sample Extraction  Buffer into a centrifugal vial.

- Add 4 mL of Acetonitrile, vortex for 3 minutes at maximum speed manually or 10 minutes using a multi-tube vortexer.

- Centrifuge for 5 minutes at 4,000 x g at room temperature (20 – 25C / 68 – 77F).

- Transfer 1.0 mL of the supernatant to another tube (corresponding to 0.5 g of the original sample) into a new vial and use a rotary evaporator to dry the sample in a 60-70C water bath under reduced pressure. Alternatively, the sample can be dried by blowing nitrogen gas in a 60-70C water bath.

- Dissolve the dried residue in 1mL of Sample Dilution Buffer .

- Centrifuge for 5 minutes at 4,000 x g at room temperature (20 – 25C / 68 – 77F).

- Use 50 L for the assay.

Note: Dilution factor: 2

### Urine

- Centrifuge 0.5 mL of the sample at 4,000 x g for 5 minutes.

- Dilute the supernatant with 1X Sample Extraction Buffer (1:4) (e. g. 100 μL of supernatant + 400 μL of buffer). Take 50 μL of the diluted sample per well for the assay.

Note: Dilution factor: 5.  To  reduce the  background,you can dilute the sample  with 1x sample dilution buffer (1:9)(e.g20 μL of supernatant + 180 μL of buffer). Dilution factor 10.

DIMETRIDAZOLE ELISA TEST KIT PROTOCOL

### Reagent Preparation

IMPORTANT: All reagents（store antibody at 4℃ before use） should be brought up to room temperature before use (1 – 2 hours at 20 – 25C / 68 – 77F); Make sure you read “Warnings and Precautions” section on page 3. Solutions should be prepared just prior to ELISA test.  All reagents should be mixed by gently inverting or swirling prior to use. Prepare volumes that are needed for the number of wells being run. Do not return the reagents to the original stock tubes/bottles. Using disposable reservoirs when handling reagents can minimize the risk of contamination and is recommended.

Preparation of 1X Wash Solution

Mix 1 volume of 20X Wash Solution with 19 volumes of distilled water.

Preparation of 1X HRP -Antibody #2

Mix 1 volume of 100x HRP Antibody#2 with 99 volumes of HRP Antibody#2 Diluent.

### ELISA Testing Protocol

Label the individual strips that will be used and aliquot reagents as the following example:

| Component | Volume per Reaction | 24 Reactions |
| --- | --- | --- |
| Antibody#1 | 100L | 2.4mL |
| 1X HRP- Antibody#2 | 150L | 3.6mL |
| 1X Wash Solution | 2.0mL | 48 mL |
| Stop Buffer | 100L | 2.4mL |
| TMB Substrate | 100L | 2.4mL |

- Add 50 L of each Standards in duplicate into different wells ( Add standards to plate only in the order from low concentration to high concentration).

- Add 50 L of each sample in duplicate into different sample wells.

- Add 100 L of Antibody#1 to each well and mix by gently rocking the plate manually for 1 minute.

- Incubate the plate for 45 minutes at room temperature (20 – 25C / 68 – 77F).

- Wash the plate 3 times with 250 L of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels ( Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

- Add 150 L of 1x HRP Antibody#2 and mix well by gently rocking the plate manually for 1 minute.

- Incubate the plate for 20 minutes at room temperature (20 – 25C / 68 – 77F).

- Wash the plate 3 times with 250 L of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels ( Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

- Add 100L of TMB substrate. Time the reaction immediately after adding the substrate. Mix the solution by gently rocking the plate manually for 1 minute while incubating.

- After incubating for 15 minutes at room temperature (20 – 25C / 68 – 77F), add 100 L of Stop Buffer to stop the enzyme reaction.

- Read the plate as soon as possible following the addition of Stop Buffer on a plate reader with 450 nm wavelength.

### Dimetridazole Concentration Calculations

A standard curve can be constructed by plotting the mean relative absorbance (%) obtained from each reference standard against its concentration in ng/mL on a logarithmic curve.

Relative absorbance (%) = 2

Use the mean relative absorbance values for each sample to determine the corresponding concentration of the tested drug in ng/mL from the standard curve.

The following figure is a typical dimetridazole standard curve.

REAGEN LLC

Add.: 102CommerceDr.Unit8,Moorestown,NJ08057

Tel.:856-727-0250          FAX.:856-727-0251

E-mail:reagenkits@gmail.com   reagenllc@gmail.com

Web.:www.reagen.us