REAGEN™

Erythromycin ELISA Test Kit

RND99052

Product Description

REAGEN™ Erythromycin ELISA Test Kit is a competitive enzyme immunoassay for the quantitative analysis of Erythromycin in water, urine,serum, meat(beef,chicken and pork),fish and shrimp samples. The unique features of the kit are:

- High recovery (80-105%),rapid(less than 40 minutes) extraction methods for various samples.

- High sensitivity (0.5ng/g or ppb)and low detection limit (2.5ppb)for fish and shrimp samples.

- High reproducibility..

- A quick ELISA assay(less than 2 hours regardless of number of samples).

Procedure Overview

The method is based on a competitive colorimetric ELISA assay. The drug of interest has been coated in the plate wells. During the analysis, sample is added along with the primary antibody specific for the target drug. If the target is present in the sample, it will compete for the antibody, thereby preventing the antibody from binding to the drug attached to the well. The secondary antibody, tagged with a peroxidase enzyme, targets the primary antibody that is complexed to the drug coated on the plate wells. The resulting color intensity, after addition of substrate, has an inverse relationship with the target concentration in the sample.

Kit Contents, Storage and Shelf Life

REAGEN™ Erythromycin ELISA Test Kit has the capacity for 96 determinations or testing of 42 samples in duplicate (assuming 12 wells for standards). Return any unused microwells to the foil bag and reseal them with the desiccant provided in the original package. Store the kit at 2-8°C *. The shelf life is 12 months when the kit is properly stored.

| Kit Contents | Amount | Storage |
| --- | --- | --- |
| Erythromycin Coated Plate | 1x 96-well Plate(8wells x12 strips) | 2-8°C |
| Erythromycin Standards: Negative control (white cap tube) 0.5ng/mL(yellow cap tube) 1.0 ng/mL(orange cap tube) 2.0 ng/mL (pink cap tube) 4.0 ng/mL(purple cap tube) 8.0 ng/mL (blue cap tube) 100 µg/mL (spiking, optional, red cap tube) | 1.0 mL 1.0 mL 1.0 mL 1.0 mL 1.0 mL 1.0 mL | 2-8°C * |
| 100 ng/mL (spiking, optional, red cap tube) | 1.0 mL |  |
| Erythromycin Antibody #1 | 12 mL | 2-8°C |
| 1X HRP Conjugate Antibody #2 | 18 mL |  |
| 20X Wash Solution | 30 mL |  |
| Stop Buffer | 14 mL |  |
| TMB Substrate | 10 mL |  |
| 10X PBS | 25 mL |  |

*  If you are not planning to use the kit for over 1 month, store  Erythromycin     Standards, Erythromycin Antibody and  1X HRP Conjugate at -20°C or in a freezer.

Sensitivity (Detection Limit)

| Sample Type | Detection Limit (ppb) |
| --- | --- |
| Dirty Water | 2.5 |
| Meat/Fish/shrimp | 2.5 |
| Urine/Serum | 2.5 |

Specificity (Cross-Reactivity)

| Analytes | Cross-Reactivity (%) |
| --- | --- |
| Erythromycin | 100 |
| Erythromycin Thiocyanate | 110 |
| Erythromycin Ethylsuccinate | 68 |
| Kanamycin | < 0.1 |
| Tilmicosin | < 0.1 |
| Tylosin | < 0.1 |

Required Materials Not Provided With the Kit

- Microtiter plate reader (450 nm)

- Tissue Mixer (e.g. Omni TissueMaster Homogenizer)

- Vortex mixer (e.g. Genie Vortex mixer from VWR)

- 10, 20, 100 and 1000 µL pipettes

- Multi-channel pipette: 50-300 µL (Optional)

Warnings and Precautions

- The standards contain Erythromycin. Handle with particular care.

- Do not use the kit past the expiration date.

- Do not intermix reagents from different kits or lots except for components with the same part No’s within  their  expiration  dates. ANTIBODIES  AND PLATES  ARE  KIT-AND LOT-SPECIFIC.  Make sure that the HRP Conjugate and Diluent are mixed in correct volumes.

- Try to maintain a laboratory temperature of 20–25°C (68–77°F). Avoid running assays under or near air vents, as this may cause excessive cooling, heating and/or evaporation. Also, do not run assays  in  direct sunlight,  as this may cause excessive  heat and evaporation. Cold  bench tops should be avoided by placing several layers of paper towel or some other insulate material under the assay plates during incubation.

- Make sure you are using only distilled or deionized water since water quality is very important.

- When pipetting samples or reagents  into an empty microtiter plate, place the pipette tips in the lower corner of the well, making contact with the plastic.

- Incubations of assay plates should be timed as precisely as possible. Be consistent when adding standards to the assay plate. Add your standards  first and then your samples.

- Add standards  to plate only in the order from low concentration to high concentration, as this will minimize the risk of compromising the standard curve.

- Always  refrigerate  plates  in  sealed  bags  with  a  desiccant  to  maintain  stability.  Prevent condensation from forming on plates by allowing them equilibrate to room temperature  (20 – 25°C /68 – 77°F) while in the packaging.

SAMPLE PREPARATION

Be sure samples are properly stored. In general, samples should be refrigerated at 2-4°C for no more than 1-2 days. Freeze samples to a minimum of -20°C  if  they need to be stored for a longer period. Frozen samples can be thawed at room temps (20 – 25°C / 68 – 77°F) or in a refrigerator before use.

Preparation of 1X PBS

Mix 1 volume of the 10 X PBS with 9 volumes of distilled water.

### Dirty Water

- Acquire 2 mL dirty water sample, adjust pH to 6~7 with minimum amount of acid or base.

- Take 1.0 mL of the above dirty water sample and add 4 mL of 35% ethanol/65% 1X PBS.

- Vortex at maximum speed for 3 minutes manually or vortex vigorously for 10 minutes in a multi-tube vortexer.

- Centrifuge for 5 minutes at 4,000 x g at room temperature (20–25°C (68–77°F). to obtain a clean supernatant. Alternatively, filter through a filtration paper to obtain a clean supernatant.

- Use 50 μL of the supernatant for the assay.

Note: Dilution factor: 5.

### Meat/Fish /shrimp

- Remove fat from the sample. Homogenize the sample with a suitable mixer.

- Weigh out 1.0 g of the homogenized sample and add 4 mL of 35% ethanol/65% 1X PBS.

- Vortex for 15 minutes at maximum speed in a multi-tube vortex mixer or vortex vigorously for 3 minutes manually.

- Centrifuge for 10 minutes at 4,000 x g at room temperature (20–25°C (68–77°F).

- Transfer 0.75 mL of the supernatant to a tube, add 0.75 mL of hexane. Vortex vigorously for 1 minute.

- Centrifuge for 4 minutes at 4,000 x g at room temperature (20–25°C (68–77°F).

Remove top hexane layer completely. Use 50 μL bottom layer for the assay.

Note: Dilution factor: 5.

### Urine/Serum

- Take 1.0 mL of urine sample and add 4 mL of 35% ethanol/65% 1X PBS.

- Vortex at maximum speed for 30 seconds vigorously.

- Centrifuge for 5 minutes at 4,000 x g at room temperature (20–25°C (68–77°F) to obtain a clean supernatant.

- Use 50 μL of the supernatant for the assay.

Note: Dilution factor: 5.

ERYTHROMYCIN ELISA TEST KIT PROTOCOL

Reagent Preparation

IMPORTANT: All reagents should be brought up to room temperature before use (1 – 2 hours at20 – 25°C / 68 – 77°F); Make sure you read “Warnings and Precautions” section on page 3. Solutions should be prepared just prior to ELISA test. All reagents should be mixed by gently inverting or swirling prior to use. Prepare only the volumes that are needed for the number of wells being run. Do  not return the  reagents  to the  original stock  tubes/bottles. It is recommended  that disposable reservoirs be used when handling reagents to minimize the risk of contamination.

Preparation of 1X Wash Solution

Mix 1 volume of the 20 X Wash Solution  with 19 volumes of distilled water.

ELISA Testing Protocol

Label the individual strips that will be used and aliquot reagents as the following example:

| Component | Volume per Reaction | 24 Reactions |
| --- | --- | --- |
| Erythromycin Antibody #1 | 100 µL | 2.4 mL |
| 1X HRP Conjugate #2 | 150 µL | 3.6 mL |
| 1X Wash Solution | 2.0 mL | 48 mL |
| Stop Buffer | 100 µL | 2.4 mL |
| TMB Substrate | 100 µL | 2.4 mL |

- Add 50 µL of each Erythromycin Standard in duplicate into different wells (Add standards to plate only in the order from low concentration to high concentration).

- Add 50 µL of each sample in duplicate into different sample wells.

- Add 100 µL of Antibody#1 and mix well by gently rocking the plate for 1 minute.

- Incubate the plate for 30 minutes at room temperature (20–25°C / 68–77°F).

- Wash the plate 3 times with 250 µL of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels ( Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

- Add 150 µL of 1X Antibody #2 to each well. Incubate the plate for 30 minutes at room temperature (20–25°C (68–77°F) (Avoid direct sunlight and cold bench tops during the incubation. Covering the microtiter plate while incubating is recommended).

- Wash the plate 3 times with 250 μL of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels (Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

- Add 100 µL of TMB substrate. Time the reaction immediately after adding the substrate. Mix the  solution by gently rocking the plate manually for 1 minute while incubating at room  temperature.

- After incubating for 20 minutes at room temperature (20–25°C / 68–77°F), add 100 µL of  Stop Buffer to stop the enzyme reaction.

- Read the plate as soon as possible following the addition of Stop Buffer on a plate reader with 450 nm wavelength.

Erythromycin Concentration Calculations

A standard  curve can be constructed by plotting the mean relative absorbance  (%) obtained from each reference standard against its concentration in ng/mL on a logarithmic curve.

Relative absorbance (%) =

Use the mean relative absorbance values for each sample to determine the corresponding concentration of  the tested  drug in ng/mL from the  standard  curve.

The following figure is a typical Erythromycin  standard curve

REAGEN LLC

Add.: 102CommerceDr.Unit8,Moorestown,NJ08057

Tel.:856-727-0250         FAX.:856-727-0251

E-mail:reagenkits@gmail.com  reagenllc@gmail.com

Web.: www.reagen.us