REAGEN™

Doxycycline ELISA Test Kit

RND99048

### Product Description

REAGEN™ Doxycycline ELISA Test Kit is a competitive enzyme immunoassay for the quantitative analysis of doxycycline in honey, milk,egg, urine,serum, fish, shrimp and meat. Doxycycline is a broad spectrum antibiotic which is frequently employed in animal production for its excellent antibacterial and pharmacokinetic properties.

The REAGEN™ Doxycycline ELISA Test Kit enables international and government regulatory agencies, food manufacturers and processors, as well as quality assurance organizations, to detect doxycycline in honey, milk or meat and to satisfy customer concerns about food safety. The unique features of the kit are:

- High recovery rate

- Rapid & cost-effective extraction method

- High sensitivity (0.15 ng/g or ppb) and low detection limit for samples

- Quick ELISA assay (less than 2 hours regardless of number of samples)

- High reproducibility

### Procedure Overview

The method is based on a competitive colorimetric ELISA assay. The drug of interest has been coated in the plate wells. During the analysis, sample is added along with the primary antibody specific for the target drug. If the target is present in the sample, it will compete for the antibody, thereby preventing the antibody from binding to the drug attached to the well. The secondary antibody, tagged with a peroxidase enzyme, targets the primary antibody that is complexed to the drug coated on the plate wells. The resulting color intensity, after addition of substrate, has an inverse relationship with the target concentration in the sample.

### Kit Contents, Storage and Shelf Life

The Doxycycline ELISA Test Kit has the capacity for 96 determinations or testing of 42 samples in duplicate (assuming 12 wells for standards). Return any unused microwells to the foil bag and reseal them with the desiccant provided in the original package. Store the kit at 2 – 8C *. The shelf life is 12 months when the kit is properly stored.

| Kit Contents | Amount | Storage |
| --- | --- | --- |
| Doxycycline-coated Plate | 1 x 96-well Plate (8 wells x 12 strips) | 2 - 8C |
| Doxycycline Standard Stock (450 ng Powder) Empty Vials for Standards: Negative control (white CAP tube) 0.15 ng/mL (yellow CAP tube) 0.375 ng/mL (orange CAP tube) 0.75 ng/mL (pink CAP tube) 1.5 ng/mL (purple CAP tube) 4.5 ng/mL (blue CAP tube) | 3 1 1 1 1 1 1 | -20C 2 - 8C |
| Doxycycline Antibody #1 | 12 mL | 2 - 8C * |
| 100X HRP-Conjugated Antibody #2 | 0.3 mL | 2 - 8C * |
| Antibody #2 Diluent ** | 20 mL | 2 - 8C |
| 20X Wash Solution ** | 28 mL | 2 - 8C |
| Stop Buffer ** | 14 mL | 2 - 8C |
| TMB Substrate** | 12 mL | 2 - 8C |
| Standard Diluent | 28 mL | 2 - 8C |
| 5X OXYTET Extraction Buffer (Optional) | 2 x 25 mL | 2 - 8C |
| 10X TET Sample Diluent (Optional) | 28 mL | 2 - 8C |
| TET Balance Buffer Concentrate (Optional) | 5 mL | 2 - 8C |

* If you are not planning to use the kit for over 1 month, storing Doxycycline Antibody #1 at 4 C and 100X HRP-Conjugated Antibody #2 at -20 C is recommended.

### Sensitivity (Detection Limit)

| Sample Type | Detection Limit (ng/g or ppb) |
| --- | --- |
| Egg/Cheese/Sour Cream | 1.5 |
| Honey | 3 |
| Meat/Meat Products/Fish/Shrimp/Butter | 1.5 |
| Milk/Soured Milk/Yogurt/Curd | 1.5 |
| Reconstituted Milk Powder | 1.5 |
| Reconstituted Whey Powder | 1.5 |
| Serum | 6 |
| Urine | 6 |

### Specificity (Cross-Reactivity)

| Analytes | Cross-Reactivity (%) |
| --- | --- |
| Doxycycline | 100.0 |
| Oxytetracycline | >100 |
| Tetracycline | >100 |
| Chlortetracycline | >100 |

### Required Materials Not Provided With the Kit

- Microtiter plate reader (450 nm primary filter, 630 nm differential filter)

- Vortex mixer (e.g. Genie Vortex mixer from VWR)

- 10, 20, 100 and 1000 uL pipettes

- Multi-channel pipette: 50-300 uL (Optional)

- n-hexane

- 1 M NaHCO3 in deionized water

- 1X PBS: To make 1 L, dissolve 0.24 g KH2PO4 + 1.44 g Na2HPO4 + 8 g NaCl, + 0.2 g KCl in deionized water to a final volume of 1 L.  (Ask your local distributor about ordering a pre-made 10X PBS solution.)

### Warnings and Precautions

- The standards contain Doxycycline. Handle with particular care.

- Do not use the kit past the expiration date.

- Do not intermix reagents from different kits or lots except for components with the same part No’s within their expiration dates. ANTIBODIES AND PLATES ARE KIT- AND LOT- SPECIFIC.

- Try to maintain a laboratory temperature of 20° – 25 °C (68° – 77 °F). Avoid running assays under or near air vents, as this may cause excessive cooling, heating and/or evaporation. Also, do not run assays in direct sunlight, as this may cause excessive heat and evaporation. Cold bench tops should be avoided by placing several layers of paper towel or some other insulation material under the assay plates during incubation.

- Make sure you are using only distilled or deionized water since water quality is very important.

- When pipetting samples or reagents into an empty microtiter plate, place the pipette tips in the lower corner of the well, making contact with the plastic.

- Incubations of assay plates should be timed as precisely as possible. Be consistent when adding standards to the assay plate. Add your standards first and then your samples.

- Add standards to plate only in the order from low concentration to high concentration as this will minimize the risk of compromising the standard curve.

- Always refrigerate plates in sealed bags with a desiccant to maintain stability. Prevent condensation from forming on plates by allowing them equilibrate to room temperature (20 – 25 C / 68 – 77 F) while in the packaging.

# SAMPLE PREPARATION

Be sure samples are properly stored. In general, samples should be refrigerated at 2 – 4°C for no more than 1 – 2 days. Freeze samples to a minimum of -20°C if they need to be stored for a longer period. Frozen samples can be thawed at room temps (20 – 25C / 68 – 77F) or in a refrigerator before use.

- Preparation of 1X OXYTET Extraction Buffer

Mix 1 volume of 5X OXYTET Extraction Buffer with 4 volumes of distilled water.

2.  Preparation of 1X TET Sample Diluent

Mix 1 volume of 10X TET Sample Diluent with 9 volumes of distilled water.

3.  Preparation of 1X OXYTET Sample Balance Buffer

Combine 1.5 mL of TET Balance Buffer Concentrate and 1.5 mL of deionized water.

### Egg/Cheese/Sour Cream/Milk/Soured Milk/Yogurt

- To 1 g of sample (or 1 mL of liquid sample), add 3 mL of 1X OXYTET Extraction Buffer.

- Vortex for 10 minutes in a multi-tube vortexer or shake 30 minutes on a shaker.

- Centrifuge for 10 minutes at 4,000 x g.

- Transfer 200 uL of the supernatant to a new tube containing 25 uL of 1X OXYTET Sample Balance Buffer. Swirl the tube. Then add 275 uL of 1X TET Sample Diluent.  Note: When transferring the supernatant, be sure to avoid any thick layer that may be sitting on top of the sample.

- Vortex for 1 minute.

- Use 100 uL per well in the assay.

Note: Dilution factor: 10

When performing analysis of samples, if you observe any false negatives (B/Bo values higher than 1.0), or higher than normal recovery rates, this may be because the starting sample is too acidic.  In this case, for Step 1 above, add 50 uL of 1 M NaHCO3 (aq.) solution to the initial sample, swirl to dissolve, then add the 1X OXYTET Extraction Buffer and continue with the procedure.

### Honey

1.  Weigh 0.5 g of honey into a centrifuge vial.

2.  Add 3.5 mL of 1X PBS.

3.  Ensure the sample is fully dissolved in the PBS by heating for 5 – 10 minutes at 50 °C.

4.  Vortex for 10 minutes in a multi-tube vortexer or shake 30 minutes on a shaker.

5.  Centrifuge for 10 minutes at 4,000 x g.

6.  Transfer 200 uL of the supernatant to a new tube containing 300 uL of 1X PBS.

7.  Vortex for 1 minute.

8.  Use 100 uL per well in the assay.

Note: Dilution factor: 20

### Meat/Meat Products/Fish/Shrimp/Butter

- To 1 g of homogenized sample, add 3 mL of 1X OXYTET Extraction Buffer and 1 mL n-hexane. For shrimp samples ONLY: DO NOT add n-hexane.

- Vortex for 10 minutes in a multi-tube vortexer or shake 30 minutes on a shaker.

- Centrifuge for 10 minute at 4,000 x g.

- Discard the upper n-hexane layer.

- Transfer 200 uL of the supernatant to a new tube containing 25 uL of 1X OXYTET Sample Balance Buffer. Swirl the tube. Then add 275 uL of 1X TET Sample Diluent. Note: When transferring the supernatant, avoid any hexanes and any thick layer or interphase that may be sitting on top of the sample.

- Vortex for 1 minute.

- Use 100 uL per well in the assay.

Note: Dilution factor: 10

To detect Doxycycline in the range of 15 – 450 ppb: After Step 5, transfer 100 uL of this sample to 900 uL of 1X TET Sample Diluent.  Vortex for 1 minute. The dilution factor will be 100.  Proceed to ELISA protocol.

### Milk Powder

- Reconstitute 1 g of dry milk powder with 9 mL of deionized or distilled water. Mix well.

- To 1 mL of sample, add 3 mL of 1X OXYTET Extraction Buffer.

- Vortex for 10 minutes in a multi-tube vortexer or shake 30 minutes on a shaker.

- Centrifuge for 10 minutes at 4,000 x g.

- Transfer 200 uL of the supernatant to a new tube containing 25 uL of 1X OXYTET Sample Balance Buffer. Swirl the tube. Then add 275 uL of 1X TET Sample Diluent.

- Vortex for 1 minute.

- Use 100 uL per well in the assay.

Note: Dilution factor: 10

### Whey Powder

- Reconstitute 1 g of dry whey powder with 12 mL of deionized or distilled water. Mix well.

- To 1 mL of sample, add 75 uL of 1 M NaHCO3 (aq). Swirl the tube, then add 3 mL of 1X OXYTET Extraction Buffer.

- Vortex for 10 minutes in a multi-tube vortexer or shake 30 minutes on a shaker.

- Centrifuge for 10 minutes at 4,000 x g.

- Transfer 200 uL of the supernatant to a new tube containing 25 uL of 1X OXYTET Sample Balance Buffer. Swirl the tube. Then add 275 uL of 1X TET Sample Diluent.

- Vortex for 1 minute.

- Use 100 uL per well in the assay.

Note: Dilution factor: 10

### Serum

- Take 0.2ml of the sample, centrifuge at 4000g for 5 min.

- Take 0.1ml of the supment, add 3.9ml of 10mm PBS buffer ,PH7.4

- Vortex for 1 min.

- Use 75 uL per well for the assay.

Note: Dilution factor:40

### Urine

- Centrifuge 1.5ml of the urine sample at 4,000 x g for 5 minutes.

- Take 0.5ml of the supment, add 19.5ml of 10mm PBS buffer, PH 7.4

- Vortex for 1 min.

- Use 75 uL per well for the assay.

Note: Dilution factor: 40

# DOXYCYCLINE ELISA TEST KIT PROTOCOL

### Reagent Preparation

IMPORTANT: All reagents should be brought up to room temperature before use (1 – 2 hours at 20 – 25C / 68 – 77F); Make sure you read “Warnings and Precautions” section. Solutions should be prepared just prior to ELISA test. All reagents should be mixed by gently  inverting or swirling prior to use. Prepare volumes that are needed for the number of wells being run. Do not return the reagents to the original stock tubes/bottles. Using disposable reservoirs when handling reagents can minimize the risk of contamination and is recommended.

- Preparation of Doxycycline Working Standards

The Doxycycline standard is provided as a 450 ng lyophilized powder. To this powder, add 1.5 mL of Standard Diluent to obtain a 300 ppb solution. Vortex the solution vigorously for 2 minutes and leave it at room temperature for 10 minutes. Repeat this procedure three times to ensure that all of the powder has dissolved. Transfer the solution to a dark plastic vial and store it at -20 °C for any future use. To make the working standards, serially dilute the 300 ppb stock solution with Standard Diluent. Vortex each standard for at least 30 seconds before continuing to the next dilution.

| Working Standards | Doxycycline Source | Volume of Source Doxycycline | Volume of Standard Diluent |
| --- | --- | --- | --- |
| 4.5 ppb vial | 300 ppb stock vial | 30 uL | 1970 uL |
| 1.5 ppb vial | 4.5 ppb vial | 500 uL | 1000 uL |
| 0.75ppb vial | 1.5 ppb vial | 500 uL | 500 uL |
| 0.375 ppb vial | 0.75 ppb vial | 500 uL | 500 uL |
| 0.15 ppb vial | 0.375 ppb vial | 500 uL | 750uL |
| Negative Control vial | N/A | 0 uL | 500 uL |

Note: The 300 ppb stock solution and the working standards must be freshly prepared before the ELISA test and these standards should be used within the same day.  When finished, empty the standard vials. These empty vials can be re-used for the next preparation of the working standards.

- Preparation of 1X Wash Solution

Mix 1 volume of the 20X Wash Solution with 19 volumes of distilled water.

- Preparation of 1X HRP-Conjugated Antibody #2

Mix 1 volume of the 100X HRP-Conjugated Antibody #2 with 99 volumes of Ab#2 Diluent.

- Tips for Optimal ELISA Performance

a) Allow the entire kit to equilibrate at room temperature for at least two hours before starting any ELISA assay.

b) Avoid light as much as possible during sample preparation and ELISA assay.

c) For plate washing steps: after addition of 250 uL wash buffer to the wells, incubate the plate for 20 – 30 seconds; shake the plate gently before pouring out the wash buffer. Repeat this procedure for each of the three washes.

d) Pipette all reagents and samples very accurately, especially the samples, even if you must slow down while pipetting. If the sample appears to have any precipitate, it is critical that you pipette the solution up and down several times before adding the 100 uL to each well.

### ELISA Testing Protocol

Label the individual strips that will be used and aliquot reagents as the following example:

| Component | Volume per Reaction | 24 Reactions |
| --- | --- | --- |
| Doxycycline Antibody #1 | 100 uL | 2.4 mL |
| 1X HRP-Conjugated Antibody #2 | 150 uL | 3.6 mL |
| 1X Wash Solution | 2.5 mL | 60 mL |
| Stop Buffer | 100 uL | 2.4 mL |
| TMB Substrate | 100 uL | 2.4 mL |

    - Add 100 uL of each Doxycycline Standards (Negative Control, 0.15, 0.375, 0.75. 1.5, 4.5 ppb) in duplicate into different wells. ( Add standards to plate in order from low concentration to high concentration.)

    - Add 100 uL of each sample in duplicate into different sample wells.  **Because of the nature of some samples, a precipitate may form after sample preparation. It is therefore required to vortex the sample thoroughly before adding to each well.

    - Add 100 uL of Antibody #1 and mix well by gently rocking the plate manually for 1 minute.

    -  Incubate the plate for 55 minutes at room temperature (20 – 25C / 68 – 77F). **Please see Tips for Optimal ELISA Performance in Reagent Preparation Section.

    - Thoroughly decant or aspirate solution from wells and discard the liquid. Wash the plate 3 times with 250 uL of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels. ( Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps.)

    - Add 150 uL of 1X Antibody #2 solution.  Incubate the plate for 25 minutes at room temperature (20 – 25C / 68 – 77F).  ( Avoid direct sunlight and cold bench tops during the incubation. Covering the microtiter plate while incubating is recommended.)

    - Thoroughly decant or aspirate solution from wells and discard the liquid. Wash the plate 3 times with 250 of uL 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels. ( Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps.)

    - Add 100 uL of TMB Substrate to each well.  Incubate the plate for 15 minutes at room temperature (20 – 25C / 68 – 77F).  Time the reaction immediately after adding the substrate. Mix the solution by gently rocking the plate manually for 1 minute while incubating. ( Do not put any substrate back to the original container to avoid any potential contamination. Covering the microtiter plate while incubating is recommended.)

    - After incubation, add 100 uL of Stop Buffer to stop the enzyme reaction.

    - Read the plate as soon as possible following the addition of Stop Buffer on a plate reader with 450 nm primary filter and 630 nm differential filter wavelengths. ( Before reading, use a lint-free wipe on the bottom of the plate to ensure no moisture or fingerprints interfere with the readings.)

### Doxycycline Concentration Calculations

A standard curve can be constructed by plotting the mean relative absorbance (%) obtained from each reference standard against its concentration in ng/mL on a logarithmic curve.

Relative absorbance (%) =

Use the mean relative absorbance values for each sample to determine the corresponding concentration of the tested drug in ng/mL from the standard curve. The following figure is a typical doxycycline standard curve.

REAGEN LLC

Add.: 102CommerceDr.Unit8,Moorestown,NJ08057

Tel.:856-727-0250        FAX.:856-727-0251

E-mail:reagenkits@gmail.com reagenllc@gmail.com

Web.: www.reagen.us