REAGEN™

Ethoxyquin ELISA Test Kit

RND99034

Product Description

REAGEN™ Ethoxyquin ELISA Test Kit is a competitive enzyme immunoassay for the quantitative analysis of ethoxyquin in fish, shrimp, feed and meat. Ethoxyquin is a quinoline-based antioxidant used as a food preservative and a pesticide. To date, the U.S. Food and Drug Administration (FDA) has only found a verifiable connection between ethoxyquin and buildup of protoporphyrin IX in the liver, as well as elevations in liver-related enzymes in some animals, but there are no known health consequences from these effects. In Japan, the ethoxyquin content must not be higher than 10 ppb in shrimp, 1 ppm in fish.

The REAGEN™ ethoxyquin ELISA Test Kit enables international and government   regulatory agencies, food manufacturers and processors, as well as quality assurance organizations, to detect ethoxyquin in fish, shrimp to satisfy customer’s concerns about food safety. The unique features of the kit are:

- High recovery (>75%) and cost-effective extraction methods

- High sensitivity (0.1 ng/g or ppb)

- Quick ELISA assay (less than 2 hours regardless of the number of samples)

- High reproducibility

Procedure Overview

The method is based on a competitive colorimetric ELISA assay. The ethoxyquin conjugate has been coated in the plate wells. During the analysis, a sample is added along with the primary antibody specific for ethoxyquin. If the target is present in the sample, it will compete for the antibody, thereby preventing the antibody from binding

to the ethoxyquin attached to the wells. The secondary antibody, tagged with a

peroxidase enzyme, targets the primary antibody associated with the ethoxyquin coated on the plate wells. The resulting color intensity, after addition of substrate, has an inverse relationship with the target concentration in the sample.

Kit Contents, Storage and Shelf Life

REAGEN™  Ethoxyquin ELISA Test Kit has the capacity for 96 determinations or testing of 42 samples in duplicate (assuming 12 wells for standards). Return any unused microwells to the foil bag and reseal them with the desiccant provided in the original package. Store the kit at 2-8°C. The shelf life is 12 months when the kit is properly stored.

| Kit Contents | Amount | Storage |
| --- | --- | --- |
| Ethoxyquin Coated Plate | 1 x 96-well Plate (8 wells x 12 strips) | 2-8°C |
| Ethoxyquin Standard Stock: 1000 ppb | 0.8 mL | -20°C * |
| 10X Ethoxyquin Sample Suspension Solution | 14 mL | 2-8°C * |
| Ethoxyquin Antibody #1 | 12 mL | 2-8°C * |
| 100X HRP-Conjugated Antibody #2 | 250L | 2-8°C |
| Antibody #2 Diluent | 20 mL | 2-8°C |
| Ethoxyquin Clean Up Mix | 12.5g | 2-8°C |
| 20X Wash Solution | 28 mL | 2-8°C |
| Stop Buffer | 14 mL | 2-8°C |
| TMB Substrate | 12 mL | 2-8°C |

*If you are not planning to use the kit for over 1 month, store the Ethoxyquin Antibody #1 and 100X HRP-Conjugated Antibody #2 at -20°C or in a freezer.

Sensitivity (Detection Limit)

| Sample Type | Detection Limit (ppb) |
| --- | --- |
| feed | 1 |
| Fish/Meat | 1 |
| Shrimp | 1 |

Specificity (Cross-Reactivity)

| Analytes | Cross-Reactivity (%) |
| --- | --- |
| Ethoxyquin | 100 |
| Enrofloxacin | <0.1 |
| Ciprofloxacin | <0.1 |
| Danofloxacin | <0.1 |
| Norfloxacin | <0.1 |
| Enoxacin | <0.1 |
| Pipemidic acid | <0.1 |
| Ofloxacin | <0.1 |
| Benofloxacin | <0.1 |
| Flumequin | <0.1 |

Required Materials Not Provided With the Kit

- Microtiter plate reader (450 nm)

- Incubator

- Tissue Mixer (e.g. Omni TissueMaster Homogenizer)

- Vortex mixer (e.g. Genie Vortex mixer from VWR)

- 10, 20, 100 and 1000 (L pipettes

- Multi-channel pipette: 50-300 (L (Optional)

- Evaporator (Optional)

- MgSO4.

Warnings and Precautions

- The standards contain ethoxyquin. Handle with particular care.

- Do not use the kit past the expiration date.

- Do not mix reagents from different kits or lots except for components with the same part No’s within their expiration dates. ANTIBODIES AND PLATES ARE KIT-AND LOT-SPECIFIC. Make sure that the antibody #2 and diluent are mixed in correct volumes.

- Try to maintain a laboratory temperature of 20°–25°C (68°–77°F). Avoid running assays under or near air vents, as this may cause excessive cooling, heating and/or evaporation. Also, do not run assays in direct sunlight, as this may cause excessive heat and evaporation. Cold bench tops should be avoided by placing several layers of paper towel or some other insulation material under the assay plates during incubation.

- Make sure you are using only distilled or deionized water since water quality is very important.

- When pipetting samples or reagents into an empty microtiter plate, place the pipette tips in the lower corner of the well, making contact with the plastic.

- Incubations of assay plates should be timed as precisely as possible. Be consistent when adding standards to the assay plate. Add your standards first and then your samples.

- Add standards to plate only in the order from low concentration to high concentration, as this will minimize the risk of compromising the standard curve.

- Always refrigerate plates in sealed bags with a desiccant to maintain stability. Prevent condensation from forming on plates by allowing them equilibrate to room temperature (20 – 25°C / 68 – 77°F)  while in the packaging.

SAMPLE PREPARATION

Be sure samples are properly stored. In general, samples should be refrigerated at 2-4°C for no more than 1-2 days. Freeze samples to a minimum of -20°C if they need to be stored for a longer period. Frozen samples can be thawed at room temps (20 – 25°C / 68 – 77°F) or in a refrigerator before use.

1. Preparation of 1X Ethoxyquin Sample Suspension Solution

Mix 1 volume of the 10 X Ethoxyquin Sample Suspension Solution with 9 volumes of distilled water.

Feed

- Homogenize the sample with a suitable mixer.

- Take 1.0 g of the homogenized sample and place into 15-mL glass tube in brown color or wrapped with aluminum foil.

- Add 5.0 mL of acetonitrile, vortex vigorously for 15 minutes in a multi-tube vortexer or shaker, or vortex for 3 minutes at maximum speed manually.

- Centrifuge the sample for 10 minutes at 4,000 X g at room temperature (20 – 25°C).

- Take 1.0 mL of supernatant to a new 15-mL glass tube in brown color or wrapped in aluminum foil containing 125 mg of Ethoxyquin Clean Up Mix.

- Vortex vigorously for 1 minute and leave at room temperature for 5 minutes. Centrifuge the sample for 10 minutes at 4,000 X g at room temperature (20 – 25°C / 68 – 77°F).

- Take 500 μL of supernatant to a 2-mL glass tube in brown coloror wrapped with aluminum foil. Use an evaporator to dry the sample at room temperature under reduced pressure in the dark. Alternatively, the sample can be dried by blowing nitrogen gas at room temperature. (Note: Avoid over drying the sample. If the sample is over dried and "hard crack" appears in the tube, the sample can not be resuspended well and the recovery rate could be reduced).

- Add 1 mL of 1X Ethoxyquin Sample Suspension Solution to the dried tube, vortex vigorously for 1 minute following by sonicating the sample for 5 minutes at maximum force in a water bath sonicator.

- Use 50 μL of the extraction solution per well in the ELISA test.

Note: Dilution factor: 10.

Meat/Fish

- Homogenize the sample with a suitable mixer.

- Take 2.0 g of the homogenized sample and place into 15-mL tube wrapped with aluminum foil.

- Add 6.0 mL of acetonitrile, and add 1.0 g of anhydrous MgSO4, vortex vigorously for 15 minutes in a multi-tube vortexer or shaker, or vortex for 3 minutes at maximum speed manually.

- Centrifuge the sample for 10 minutes at 4,000 X g at room temperature (20-25°C).

- Take 1.0 mL of supernatant to a new 15-mL tube wrapped in aluminum foil containing 250 mg of Ethoxyquin Clean Up Mix

- Vortex vigorously for  minute and leave at room temperature for 5 minutes. Centrifuge the sample for 10 minutes at 4,000 X g at room temperature (20-25°C).

- Take 500 μL of supernatant to a 2-mL glass tube in brown color or wrapped with aluminum foil. Use an evaporator to dry the sample at 30°C under reduced pressure in the dark. Alternatively, the sample can be dried by blowing nitrogen gas in a 30°C water bath or at room temperature for a longer time. (Note: Avoid over drying the sample. If the sample is over dried and "hard crack" appears in the tube, the sample can not be resuspended well and the recovery rate could be reduced).

- Add 625 μL of 1X Ethoxyquin Sample Suspension Solution to the dried tube and vortex vigorously for 2 minutes. Alternatively, to fully dissolve the dried residue, vortex vigorously for 1 minute following by sonicating the sample for 5 minutes at maximum force in a water bath sonicator.

- Use 50 μL of the extraction solution per well in the ELISA test.

Note: Dilution factor: 5.

Shrimp

- Homogenize the sample with a suitable mixer.

- Take 2.0 g of the homogenized sample and place into 15-mL glass tube in brown color or wrapped with aluminum foil.

- Add 6.0 mL of acetonitrile, shake the sample until all shrimp is freely floating in solvent, and add 1.0 g of anhydrous MgSO4, vortex vigorously for 15 minutes in a multi-tube vortexer or shaker, or vortex for 3 minutes at maximum speed manually.

- Centrifuge the sample for 10 minutes at 4,000 X g at room temperature (20-25°C).

- Take 1.0 mL of supernatant to a new 15-mL glass tube in brown color or wrapped in aluminum foil containing 125 mg of Ethoxyquin Clean Up Mix (If after drying, a lot of fat is left over in the tube and a high sample background is detected, instead of 125 mg, using 150 mg of Clean Up Mix is recommended).

- Vortex vigorously for 1 minute and leave at room temperature for 5 minutes. Centrifuge the sample for 10 minutes at 4,000 X g at room temperature (20-25°C/68-77°F).

- Take 500 μL of supernatant to a 2-mL glass tube in brown color or wrapped with aluminum foil. Use an evaporator to dry the sample at room temperature under reduced pressure in the dark. Alternatively, the sample can be dried by blowing nitrogen gas at room temperature. (Note: Avoid over drying the sample. If the sample is over dried and "hard crack" appears in the tube, the sample can not be resuspended well and the recovery rate could be reduced).

- Add 625 μL of 1X Ethoxyquin Sample Suspension Solution to the dried tube, vortex vigorously for 1 minute following by sonicating the sample for 5 minutes at maximum force in a water bath sonicator.

- Use 50 μL of the extraction solution per well in the ELISA test.

Note: Dilution factor:5.

ETHOXYQUIN ELISA TEST KIT PROTOCOL

Reagent Preparation

IMPORTANT: All reagents should be brought up to room temperature before use (1–2 hours at20 – 25°C / 68 – 77°F); Make sure you read “Warnings and Precautions” section on page 3. Solutions should be prepared just prior to ELISA test. All  reagents  should be mixed by gently inverting or swirling prior to use. Prepare only the volumes that are needed for the number of wells being run. Do  not return the  reagents to the original stock  tubes/bottles. It is recommended  that disposable reservoirs be used when handling reagents to minimize the risk of contamination.

1.  Preparation of 1X Wash Solution

Mix 1 volume of the 20 X Wash Solution  with 19 volumes of distilled water.

- Preparation of 1X HRP-Conjugated Antibody #2

Mix 1 volume of 100X HRP-Conjugated Antibody #2 with 99 volumes of Antibody #2 Diluent.

- Preparation of Ethoxyquin Standards

| Mix the following solutions by serial dilution starting from the highest concentration to the lowest in 2mL amber glass vials. Make standards fresh each time before performing the ELISA assay. Work Standards | Volume of Ethoxyquin Source | Volume of 1x Sample Suspension Solution |
| --- | --- | --- |
| 100 ppb | 100 µL (supplied stock solution, 1000 ppb) | 900 µL |
| 4.5 ppb | 45 µL (of 100 ppb dilution) | 955 µL |
| 2.25 ppb | 500 µL (of 4.5 ppb dilution) | 500 µL |
| 0.9 ppb | 400 µL (of 2.25 ppb dilution) | 600 µL |
| 0.3 ppb | 300 µL (of 0.9 ppb dilution) | 600 µL |
| 0.1 ppb | 300 µL (of 0.3 ppb dilution) | 600 µL |
| 0 ppb | 0 µL | 600 µL |

ELISA Testing Protocol

Label the individual strips that will be used and aliquot reagents as the following example:

| Component | Volume per Reaction | 24 Reactions |
| --- | --- | --- |
| Ethoxyquin Antibody | 100 µL | 2.4 mL |
| 1X HRP Conjugate | 150 µL | 3.6 mL |
| 1X Wash Solution | 2.0 mL | 48 mL |
| Stop Buffer | 100 µL | 2.4 mL |
| TMB Substrate | 100 µL | 2.4 mL |

- Add 50 L of each Ethoxyquin Standard in duplicate into different wells (Add standards to plate only in the order from low concentration to high concentration).

- Add 50 L of each sample in duplicate into different sample wells.

- Add 100 L of Antibody #1 and mix well by gently rocking the plate manually for 1 minute.

- Incubate the plate for 30 minutes at room temperature (20 – 25C / 68 – 77F). (Avoid direct sunlight and cold bench tops during the incubation. Covering the microtiter plate while incubating is recommended).

- Wash the plate 3 times with 250 L of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels (Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

- Add 150 L of 1X Antibody #2 solution. Incubate the plate for 30 minutes at room temperature (20 – 25C / 68 – 77F) (Avoid direct sunlight and cold bench tops during the incubation. Covering the microtiter plate while incubating is recommended).

- Wash the plate 3 times with 250 L of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels (Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

- Add 100 L of TMB substrate. Time the reaction immediately after adding the substrate. Mix the solution by gently rocking the plate manually for 1 minute while incubating (Do not put any substrate back to the original container to avoid any potential contamination. Any substrate solution exhibiting coloration is indicative of deterioration and should be discarded. Covering the microtiter plate while incubating is recommended).

- After incubating for 15 minutes at room temperature (20 – 25C / 68 – 77F), add 100 L of Stop Buffer to stop the enzyme reaction.

- Read the plate as soon as possible following the addition of Stop Buffer on a plate reader with 450 nm wavelength (Before reading, use a lint-free wipe on the bottom of the plate to ensure no moisture or fingerprints interfere with the readings).

Ethoxyquin Concentration Calculations

A standard curve can be constructed by plotting the mean relative absorbance (%) obtained from each reference standard against its concentration in ng/mL on a logarithmic curve.

Relative absorbance (%) =

Use the mean relative absorbance values for each sample to determine the corresponding concentration of the tested drug in ng/mL from the standard curve.

The following figure is a typical Ethoxyquin standard curve

REAGEN LLC

Add.: 102CommerceDr.Unit8,Moorestown,NJ08057

Tel.:856-727-0250         FAX.:856-727-0251

E-mail:reagenkits@gmail.com  reagenllc@gmail.com

Web.: www.reagen.us