REAGEN™

Ofloxacin ELISA Test Kit

RND99021

### Product Description

REAGEN™ofloxacin ELISA Test Kit enables government agencies, food manufacturers and quality assurance organizations to detect ofloxacin to as low as 0.1 ppb level in various sample types and to satisfy customer concerns about food safety. The features of the kit are:

- The extraction of ofloxacin from meat/tissue with a high recovery rate.

- A quick ELISA assay (less than 1.5 hours regardless of number of samples).

- High reproducibility.Procedure Overview

### Procedure Overview

- The method is based on a competitive colorimetric ELISA assay. The ofloxacin antibody has been coated in the plate wells. During the analysis, sample is added along with the ofloxacin-horseradish peroxidase (Ofloxacin-HRP) conjugate. If the ofloxacin residue is present in the sample, it will compete for the ofloxacin antibody, thereby preventing the Ofloxacin-HRP from binding to the antibody attached to the well. The resulting color intensity, after addition of the HRP substrate (TMB), has an inverse relationship with the ofloxacin residue concentration in the sample.

### Kit Contents, Storage and Shelf Life

REAGEN™ ofloxacin ELISA Test Kit has the capacity for 96 determinations or testing of 42 samples in duplicate (assuming 12 wells for standards). Return any unused microwells to the foil bag and reseal them with the desiccant provided in the original package. Store the kit at 2-8°C. The shelf life is 12 months when the kit is properly stored.

| Kit Contents | Amount | Storage |
| --- | --- | --- |
| 0floxacinAb-coated Plate | 1 x 96-well plate (8 wells x 12 strips) | 2-8°C |
| Ofloxacin Standards: Negative control (white cap tube) 0.1 ng/mL (yellow cap tube) 0.25 ng/mL (orange cap tube) 0.5 ng/mL (pink cap tube) 1.0 ng/mL (purple cap tube) 5.0 ng/mL (blue cap tube) 100 ng/mL (spiking, optional, red cap tube) | 0.8 mL 0.8 mL 0.8 mL 0.8 mL 0.8 mL 0.8 mL 0.8 mL | 2-8°C |
| Ofloxacin-HRP | 12 mL | 2-8°C |
| 10X Sample Extraction Buffer (Optional) | 25 mL |  |
| 20X Wash Solution ** | 28 mL |  |
| Stop Buffer ** | 14 mL |  |
| TMB Substrate ** | 12 mL |  |
| Meat Extraction Buffer I (Optional) | 5 mL |  |
| Meat Extraction Buffer II (Optional) | 5 mL |  |
| Sample Balance Buffer (Optional) | 25 mL |  |

*If you are not planning to use the kit for over 3 months, store Ofloxacin-HRP at -20°C preferably or in a freezer.

Sensitivity (Detection Limit)

| Sample Type | Detection Limit (ng/g or ppb) |
| --- | --- |
| Meat/Liver/Kidney/Fish/Shrimp | 0.1 (Organic Solvent Method); 2 (Rapid Method) |

Specificity (Cross-Reactivity)

| Analytes | Cross-Reactivity (%) |
| --- | --- |
| Ofloxacin | 100 |
| Ciprofloxacin | 82 |
| Enrofloxacin | 80 |
| Danofloxacin | 42 |
| Enoxacin | 14 |
| Pipemidic acid | 11 |
| Benofloxacin | 3 |
| Oxolin acid | 4 |
| Nalidixic acid | 2 |

### Required Materials Not Provided With the Kit

- Microtiter plate reader (450 nm)

- Incubator

- Tissue Mixer (e.g. Omni TissueMaster Homogenizer)

- Vortex mixer (e.g. Gneie Vortex mixer from VWR)

- 10, 20, 100 and 1000 L pipettes

- Multi-channel pipette: 50-300 L (Optional)

### Warnings and Precautions

- The standards contain ofloxacin. Handle with particular care.

- Do not use the kit past the expiration date.

- Do not mix reagents from different kits or lots except for components with the same part No’s within their expiration dates. HRP CONJUGATES AND PLATES ARE KIT-AND LOT-SPECIFIC.

- Try to maintain a laboratory temperature of 20–25°C (68°–77°F). Avoid running assays under or near air vents, as this may cause excessive cooling, heating and/or evaporation. Also, do not run assays in direct sunlight, as this may cause excessive heat and evaporation. Cold bench tops should be avoided by placing several layers of paper towel or some other insulation material under the assay plates during incubation.

- Make sure you are using only distilled or deionized water since water quality is very important.

- When pipetting samples or reagents into an empty microtiter plate, place the pipette tips in the lower corner of the well, making contact with the plastic.

- Incubations of assay plates should be timed as precisely as possible. Be consistent when adding standards to the assay plate. Add your standards first and then your samples.

- Add standards to plate only in the order from low concentration to high concentration as this will minimize the risk of compromising the standard curve.

- Always refrigerate plates in sealed bags with a desiccant to maintain stability. Prevent condensation from forming on plates by allowing them equilibrate to room temperature (20–25°C / 68–77°F) while in the packaging.

SAMPLE PREPARATION

Be sure samples are properly stored. In general, samples should be refrigerated at 2-4°C for no more than 1-2 days. Freeze samples to a minimum of -20°C if they need to be stored for a longer period. Frozen samples can be thawed at room temps (20 – 25C / 68 – 77F) or in a refrigerator before use.

- Preparation of Acetonitrile containing 1% Acetic Acid

Mix 99 volumes of acetonitrile with 1 volume of acetic acid.

- Preparation of 35% Methanol/Sample Balance Buffer

Mix 6.5 volumes of Sample Balance Buffer with 3.5 volumes of 100% methanol.

- Preparation of 1X Sample Extraction Buffer

Mix 1 volume of 10X Sample Extraction Buffer with 9 volumes of distilled water.

- Preparation of 35% Methanol/1X Sample Extraction Buffer

Mix 6.5 volumes of 1X Sample Extraction Buffer with 3.5 volumes of 100% methanol.

### Meat/Liver/Kidney/Fish/Shrimp

Organic Solvent Method

- Remove fat from the sample. Homogenize a reasonable amount of sample with a suitable mixer.

- Weigh out 2 g of the homogenized sample and mix with 4 mL of acetonitrile containing 1% acetic acid and 2 mL of hexane, vortex 3 min at maximum speed manually or vortex 10 minutes with multi-tube vortexer.

- Centrifuge for 5 minutes at 4,000 x g at room temperature (20 – 25C / 68 – 77F).

- Remove the top hexane layer, transfer 1.2 mL of the acetonitrile supernatant (corresponding to 0.4 g of the original sample) into a new vial and use a rotary evaporator to dry the sample in a 60-70C water bath under reduced pressure. Alternatively, the sample can be dried by blowing nitrogen gas in a 60-70C water bath.

- Dissolve the dried residue with 400 µL of 35% Methanol/Sample Balance Buffer and mix by vortexing at maximum speed for 1 minute.

- Use 25 µL of the sample per well for the assay.

Note: Dilution factor: 1

Rapid Method

- Remove fat from the sample. Homogenize the sample with a suitable mixer.

- Weigh out 1.0 g of the homogenized sample and add 4 mL of 35% Methanol/Sample Extraction Buffer and 50 µL of Meat Extraction Buffer I.

- Vortex for 10 minutes at maximum speed.

- Centrifuge for 10 minutes at 4,000 x g at room temperature (20 – 25C / 68 – 77F).

- Transfer 0.5 mL of the supernatant to a tube, add 25 µL of Meat Extraction Buffer II and 1.45 mL of 35% Methanol/Sample Extraction Buffer, vortex for 30 seconds.

- Use 25 µL of the sample for the assay.

Note: Dilution factor: 20.

OXFLAXACIN ELISA TEST KIT PROTOCOL

### Reagent Preparation

IMPORTANT: All reagents should be brought up to room temperature before use (1 – 2 hours at 20 – 25C / 68 – 77F); Make sure you read “Warnings and Precautions” section on page 3.Solutions should be prepared just prior to ELISA test.  All reagents should be mixed by gently inverting or swirling prior to use. Prepare volumes that are needed for the number of wells being run. Do not return the reagents to the original stock tubes/bottles. Using disposable reservoirs when handling reagents can minimize the risk of contamination and is recommended.

1.  Preparation of 1X Wash Solution

Mix 1 volume of the 20X Wash Solution with 19 volumes of distilled water.

### ELISA Testing Protocol

Label the individual strips that will be used and aliquot reagents as the following example:

| Component | Volume per Reaction | 24 Reactions |
| --- | --- | --- |
| Ofloxacin-HRP | 100 µL | 2.4 mL |
| 1X Wash Solution | 2.0 mL | 48 mL |
| Stop Buffer | 100 µL | 2.4 mL |
| TMB Substrate | 100 µL | 2.4 mL |

- Add 25 µL of each Ofloxacin Standards in duplicate into different wells (Add standards to plate only in the order from low concentration to high concentration).

- Add 25 µL of each sample in duplicate into different sample wells.

- Add 100 µL of Ofloxacin-HRP and mix well by gently rocking the plate manually for 1 minute.

- Incubate the plate for 45 minutes at room temperature (20 – 25C / 68 – 77F).

- Wash the plate 3 times with 250 µL 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels ( Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

- Add 100 µL of TMB substrate. Time the reaction immediately after adding the substrate. Mix the solution by gently rocking the plate manually for 1 minute while incubating

- After incubating for 15 minutes at room temperature (20 – 25C / 68 – 77F), add 100 µL of Stop Buffer to stop the enzyme reaction.

- Read the plate as soon as possible following the addition of Stop Buffer on a plate reader with 450 nm wavelength

### Ofloxacin Concentration Calculations

A standard curve can be constructed by plotting the mean relative absorbance (%) obtained from each reference standard against its concentration in ng/mL on a logarithmic curve.

Relative absorbance (%) =

The following figure is a typical ofloxacin standard curve

REAGEN LLC

Add.: 102CommerceDr.Unit8,Moorestown,NJ08057

Tel.:856-727-0250         FAX.:856-727-0251

E-mail:reagenkits@gmail.com  reagenllc@gmail.com

Web.: www.reagen.us