REAGEN™

Beta-Lactam ELISA Test Kit

RND99027

### Product Description

REAGEN™ Beta-Lactam ELISA Test Kit is a competitive enzyme immunoassay for the quantitative analysis of beta-lactams including ampicillin, penicillin G, amoxicillin, cloxacillin, oxacillin in milk, urine, tissue (liver, kidney, and muscle), serum, and plasma samples. REAGEN™ Beta-lactam ELISA Test Kit enables international and government regulatory agencies, food manufacturers and processors, as well as quality assurance organizations, to detect beta-lactam containing antibiotics to satisfy customer concerns about food safety.

The unique features of the kit are:

- High recovery (75-105%), low-cost rapid sample prep protocol (10 minutes - 1 hour)

- High sensitivity: 0.4 ng/gram = 0.4 ppb (parts per billion).

- High reproducibility (interplate and intraplate CV typically <8%)

- Rapid assay protocol (less than 2 hours to process 40 samples in duplicate)

### Procedure Overview

The method is based on a competitive colorimetric ELISA assay. The drug of interest (beta-lactam antibiotic) has been coated in the plate wells. The sample to be tested is mixed with a high-affinity beta-lactam capture protein and added to the well. If the beta-lactam target is present in the sample, it will compete for the capture protein, thereby preventing the capture protein from binding to the drug attached to the well. The secondary antibody, tagged with a peroxidase enzyme, targets the capture protein bound to the drug coated on the plate wells. The resulting color intensity, after addition of substrate, has an inverse relationship with the target concentration in the sample. The levels of beta-lactam antibiotic in the sample are determined by comparison of the color intensity of unknown samples with the standard curve, which is made using the beta-lactam standard provided in the kit.

### Kit Contents, Storage and Shelf Life

| Kit Contents | Amount | Storage |
| --- | --- | --- |
| Beta-Lactam Coated Plate | 1 x 96-well Plate (8 wells x 12 strips) | 2-8C |
| Penicillin G powder stock (1000 ng): Empty Vials for Standards: Negative control (white cap tube) 0.08 ng/mL (yellow cap tube) 0.2 ng/mL (orange cap tube) 0.4 ng/mL (pink cap tube) 0.8 ng/mL(purple cap tube) 1.2 ng/mL (blue cap tube) | 2 1 1 1 1 1 1 | 2-8C * |
| 10X Beta-Lactam Capture Protein | 1.8 mL | 2-8C * |
| TMB Substrate ** | 10 mL |  |
| 200X HRP-Conjugated Antibody #2 | 2 x 100 L |  |
| Antibody #2 Diluent ** | 20 mL |  |
| Stop Buffer ** | 20 mL |  |
| 20X Wash Solution ** | 30 mL |  |
| 10X Beta-Lactam Extraction Buffer* Milk Standard Dilution Buffer 1 | 2 x 28 mL 15 mL |  |

* 	If the kit is to be stored for over two weeks before use, store Standard Dilution Buffer1, Penicillin G powder stock, 10X Beta-Lactam Capture Protein  and 200X HRP-Conjugated Antibody #2 at -20C in a freezer that does not cycle to higher temperatures (i.e. a non-frost-free freezer). To avoid repeated freeze-thawing of the capture protein it is strongly recommended to store this component in single-use aliquots.

### Sensitivity

| Antibiotic | Milk, Meat, Tissue, Fish and Urine(ppb) | Serum and Plasma(ppb) |
| --- | --- | --- |
| Penicillin G | 0.40 | 0.80 |
| Ampicillin | 0.43 | 0.86 |
| Azlocillin | 0.40* | 0.80 |
| Pipercillin | 0.52 | 1.04 |
| Nafcillin | 0.58 | 1.16 |
| Amoxicillin | 0.62 | 1.24 |
| Cloxacillin | 0.64 | 1.28 |
| Oxacillin | 0.64 | 1.28 |
| Cefoperazone | 0.67 | 1.34 |
| Ceftiofur | 0.7 | 1.4 |

### Specificity

| Analytes | Cross-Reactivity (%) |
| --- | --- |
| Penicillin G | 100 |
| Ampicillin | 92 |
| Azlocillin | 112 |
| Piperacillin | 70 |
| Nafcillin | 55 |
| Amoxicillin | 44 |
| Cloxacillin | 40 |
| Oxacillin | 40 |
| Cefoperazone | 34 |
| Ceftiofur | 25 |

### Required Materials Not Provided With the Kit

- Microtiter plate reader (450 nm)

- Incubator

- Tissue Mixer (e.g. Omni TissueMaster Homogenizer)

- Vortex mixer (e.g. Gneie Vortex mixer from VWR)

- 10, 20, 100 and 1000 L pipettes

- Multi-channel pipette: 50-300 L (Optional)

- n-Hexane

- Distilled deionized water

- Disposable plastic reservoirs

- Paper towels

### Warnings and Precautions

- The standards contain Penicillin G. Handle with particular care.

- Do not use the kit past the expiration date.

- Do not intermix reagents from different kits or lots except for components with the same part No’s within their expiration dates. ANTIBODIES AND PLATES ARE KIT-AND LOT-SPECIFIC. Make sure that the antibody #2 and diluent are mixed in correct volumes.

- Try to maintain a laboratory temperature of 20°–25°C (68°–77°F). Avoid running assays under or near air vents, as this may cause excessive cooling, heating and/or evaporation. Also, do not run assays in direct sunlight, as this may cause excessive heat and evaporation. Cold bench tops should be avoided by placing several layers of paper towel or some other insulation material under the assay plates during incubation.

- Make sure you are using only distilled or deionized water since water quality is very important.

- When pipetting samples or reagents into an empty microtiter plate, place the pipette tips in the lower corner of the well, making contact with the plastic.

- Incubations of assay plates should be timed as precisely as possible. Be consistent when adding standards to the assay plate. Add your standards first and then your samples.

- Add standards to plate only in the order from low concentration to high concentration as this will minimize the risk of compromising the standard curve.

- Always refrigerate plates in sealed bags with a desiccant to maintain stability. Prevent condensation from forming on plates by allowing them equilibrate to room temperature (20 – 25C / 68 – 77F) while in the packaging.

SAMPLE PREPARATION

Be sure samples are properly stored. In general, samples should be refrigerated at 2-4°C for no more than 1-2 days. Freeze samples to a minimum of -20°C if they need to be stored for a longer period. Frozen samples can be thawed at room temps (20 – 25C / 68 – 77F) or in a refrigerator before use.

- Preparation of 1X Beta-Lactam Extraction Buffer:

Mix 1 volume of 10X Beta-Lactam Buffer with 9 volumes of distilled water.

### Meat/Fish Tissue

    - Homogenize a reasonable amount of sample with a suitable mixer until the sample has a consistency like paste. Store sample at -20°C until analyzed (i.e. a non-frost-free freezer).

    - Weigh out 1 g of the homogenized sample and mix with 9 mL of 1X Beta-Lactam Extraction Buffer then vortex for 3 minutes or shake for 20 minutes with a shaker.

    - Centrifuge for 10 minutes at 4,000 x g at room temperature (20 – 25C / 68 – 77F).

    - Take out 1 mL of supernatant and mix it with 2 mL of hexane, vortex for 2 minutes.

    - Centrifuge sample for 10 minutes at 4,000 rpm.

    - Discard the top hexane layer and use 50 µL of the low aqueous layer per well in the test. Be sure to remove all of the upper layer by pipetting it and the miscible layer off before taking sample.

Note: The dilution factor in above examples is 10. If the beta-lactam concentration is expected to be high, the sample can be diluted further with Beta-Lactam Extraction Buffer.

### Milk

For fat-free milk, dilute the milk sample 1:9 with 1X Beta-Lactam Extraction Buffer (e.g. 20 L of milk + 180 L of Beta-Lactam Extraction Buffer). Use 50 L of the diluted sample per well for the assay.

For regular milk with fat, centrifuge the milk sample at 4,000 rpm for 10 minutes and discard the upper fat layer. A toothpick, pipette tip, or weighing spatula is generally adequate to drag the fat layer to the side or pierce an opening through which you may obtain the separated milk. Dilute the milk sample 1:9 with 1X Beta-Lactam Extraction Buffer (e.g. 20 L of milk + 180 L of Beta-Lactam Extraction Buffer). Use 50 L of the diluted sample per well for the assay.

Note: Dilution factor:10.

### Serum/Plasma

    - Centrifuge the serum/plasma sample at 4,000 rpm for 15 minutes.

    - Dilute the supernatant 1:19 with 1X Beta-Lactam Extraction Buffer (e.g. 10 L of serum/plasma + 190 L of Buffer).

    - Use 50 L of the diluted sample per well for the assay.

Note: Dilution factor: 20.

### Urine

    - Centrifuge the urine sample at 4,000 rpm for 10 minutes.

    - Dilute the supernatant 1:9 with 1X Beta-Lactam Extraction Buffer (e.g. 20 L of urine + 180 L of Buffer).

Note: Dilution factor: 10.

BETA-LACTAM ELISA TEST KIT PROTOCOL

### Reagent Preparation

IMPORTANT: All reagents should be brought up to room temperature before use (1 – 2 hours at 20 – 25C / 68 – 77F); Make sure you read “Warnings and Precautions” section on page 3. Solutions should be prepared just prior to ELISA test.  All reagents should be mixed by gently inverting or swirling prior to use. Prepare volumes that are needed for the number of wells being run. Do not return the reagents to the original stock tubes/bottles. Using disposable reservoirs when handling reagents can minimize the risk of contamination and is recommended.

- Preparation of Penicillin G Standards

Add 1 mL of 1X Beta-Lactam Extraction Buffer to the stock vial (1000 ng) and mix by vortexing for ~30 sec to obtain 1000 ppb stock.  Centrifuge the tube briefly to re-collect the liquid. From the stock 1000ppb tube take 20uL and dilute it into 980uL of 1X Beta-Lactam Extraction Buffer to make a 20ppb stock. To make the working standards, serially dilute the 20 ppb standard stock vial in the provided empty standard vials as described in the following table. Make sure to mix each standard vial by vortexing thoroughly and briefly centrifuge before removing aliquots.

*For Preparing Standards for use with milk: Use the Milk Standard Dilution Buffer 1 for all dilutions after the 20ppb stock provided with the kit, instead of 1X Beta-Lactam Extraction Buffer.

| Standards | Penicillin G Source | Volume of Penicillin G Solution to add from Source Tube | Volume of Diluent: (1X Beta-Lactam Extraction Buffer or Milk Standard Dilution Buffer 1) |
| --- | --- | --- | --- |
| 1.2 ppb vial | 20 ppb stock vial | 60 L | 940 L |
| 0.8 ppb vial | 1.2 ppb vial | 500 L | 250 L |
| 0.4 ppb vial | 0.8 ppb vial | 500 L | 500 L |
| 0.2 ppb vial | 0.4 ppb vial | 500 L | 500 L |
| 0.08ppb vial | 0.2ppb vial | 500 L | 750 L |
| Negative Control vial | N/A | 0 L | 500 L |

Note: The 1000 ppb stock vial can be divided into single-use aliquots and stored at -20C in a non-frost-free freezer for at least 1 month. The working standards must be freshly prepared before use in the ELISA. After using the standards, empty the standard vials and store them at 4C. The empty vials can be re-used for preparing standards for subsequent assays.

- Preparation of 1X Beta-Lactam Capture Protein

Mix 1 volume of 10X Beta-Lactam Capture Protein with 9 volumes of 1X Beta-Lactam Extraction Buffer.

- Preparation of 1X HRP-Conjugated Antibody #2

Mix 1 volume of 100X HRP-Conjugated Antibody #2 with 199 volumes of Antibody #2 Diluent.

- Preparation of 1X Wash Solution

Mix 1 volume of the 20X Wash Solution with 19 volumes of distilled water.

### ELISA Testing Protocol

Label the individual strips that will be used and aliquot reagents as the following example:

| Component | Volume per Reaction | 24 Reactions |
| --- | --- | --- |
| 1X Beta-Lactam Capture Protein | 100 L | 2.4 mL |
| 1X HRP-Conjugated Antibody #2 | 150 L | 3.6 mL |
| 1X Wash Solution | 2.0 mL | 48 mL |
| Stop Buffer | 100 L | 2.4 mL |
| TMB Substrate | 100 L | 2.4 mL |

    - Add 50 L of each Penicillin G Standards in duplicate into different wells ( Add standards to plate only in the order from low concentration to high concentration).

    - Add 50 L of each sample in duplicate into different sample wells.

    - Add 100 L of 1X Beta-Lactam Capture Protein and mix well by gently rocking the plate manually for 1 minute.

    - Incubate the plate for 30 minutes at room temperature (20 – 25C / 68 – 77F). 

    - Wash the plate 4 times with 250 L of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels ( Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

    - Add 150 L of 1X Antibody #2 solution. Incubate the plate for 30 minutes at room temperature (20 – 25C / 68 – 77F) (  Avoid direct sunlight and cold bench tops during the incubation. Covering the microtiter plate while incubating is recommended).

    - Wash the plate 4 times with 250 of L 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels ( Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

    - Add 100 L of TMB substrate. Time the reaction immediately after adding the substrate. Mix the solution by gently rocking the plate manually for 1 minute while incubating (  Do not put any substrate back to the original container to avoid any potential contamination. Any substrate solution exhibiting coloration is indicative of deterioration and should be discarded. Covering the microtiter plate while incubating is recommended).

    - After incubating for 10-15 minutes at room temperature (20 – 25C / 68 – 77F), add 100 L of Stop Buffer to stop the enzyme reaction.

    - Read the plate as soon as possible following the addition of Stop Buffer on a plate reader with 450 nm wavelength ( Before reading, use a lint-free wipe on the bottom of the plate to ensure no moisture or fingerprints interfere with the readings).

### Beta-Lactam Concentration Calculations

A standard curve can be constructed by plotting the mean relative absorbance (%) obtained from each reference standard against its concentration in ng/mL on a logarithmic curve.

Relative absorbance (%) =

Use the mean relative absorbance values for each sample to determine the corresponding concentration of the tested drug in ng/mL from the standard curve. A special program with Excel functionality, MaxSignal® ELISA Analysis Program in Excel, is available upon request to evaluate the MaxSignal® ELISA test results. Please contact your local distributor or techsupport@biooscientific.com for further information.

The following figure is a typical beta-Lactam standard curve. The sample detection and quantification limit for this kit are calculated as below.

Sample detection limit = (0.08 ng/g or ppb) x (dilution factor)

Sample quantification limit = (0.2 ng/g or ppb) x (dilution factor)

For example, the dilution factor for milk sample is 10, therefore, the detection limit for feed sample is 0.8 ng/g or ppb (0.08 ng/g x dilution factor of 10) and the quantification limit is 2 ng/g or ppb (0.2 ng/g x dilution factor 10).

REAGEN LLC

Add.: 102CommerceDr.Unit8,Moorestown,NJ08057

Tel.:856-727-0250         FAX.:856-727-0251

E-mail:reagenkits@gmail.com  reagenllc@gmail.com

Web.: www.reagen.us