REAGEN™

Tetracycline(TET) ELISA Test Kit

RND99025

### Product Description

REAGEN™ Tetracycline (TET) ELISA Test Kit is a competitive enzyme immunoassay for the quantitative analysis of Tetracycline in in honey, butter, whey, egg, cheese, milk, meat, meat products such as sausage, fish and shrimp.

The unique features of the kit are:

- High recovery (75-120%), rapid (less than 30 minutes), and cost-effective extraction methods.

- High sensitivity (0.05 ng/g or ppb) and low detection limit for samples.

- High reproducibility.

- A quick ELISA assay (less than 2 hours regardless of number of samples).

### Procedure Overview

REAGEN™ Tetracycline (TET) ELISA Test Kit is based on a competitive colorimetric ELISA assay. The drug of interest has been coated in the plate wells. During the analysis, sample is added along with the primary antibody specific for the target drug. If the target is present in the sample, it will compete for the antibody, thereby preventing the antibody from binding to the drug attached to the well. The secondary antibody, tagged with a peroxidase enzyme, targets the primary antibody that is complexed to the drug coated on the plate wells. The resulting color intensity, after addition of substrate, has an inverse relationship with the target concentration in the sample.

### Kit Contents, Storage and Shelf Life

REAGEN™ Tetracycline (TET) ELISA Test Kit has the capacity for 96 determinations or testing of 42 samples in duplicate (assuming 12 wells for standards). Return any unused microwells to the foil bag and reseal them with the desiccant provided in the original package. Store the kit at 2-8C . The shelf life is 12 months when the kit is properly stored.

| Kit Contents | Amount | Storage |
| --- | --- | --- |
| Tetracycline-Coated Plate | 1 x 96-well plate (8 wells x 12 strips) | 2-8℃ |
| Tetracycline Standards: Negative control (white Cap tube) 0.05 ng/mL (yellow Cap tube) 0.15 ng/mL (orange Cap tube) 0.4 ng/mL (pink Cap tube) 0.8 ng/mL (purple Cap tube) 1.6 ng/mL (blue Cap tube) 100ng/mL (Spikingl red Cap tube) | 1.2 mL 1.2 mL 1.2 mL 1.2 mL 1.2 mL 1.2 mL 1.2 mL | 2-8℃ |
| Tetracycline Antibody #1 | 12mL | 2-8℃ |
| 100X HRP-Conjugated Antibody #2 | 300μL |  |
| Antibody #2 Diluent | 20mL |  |
| 20X Wash Solution | 28mL |  |
| Stop Buffer | 14mL |  |
| TMB Substrate | 12mL |  |
| 5 X OXYTET Extraction Buffer(Optional) | 2X 25 mL |  |
| 10X TET Sample Diluent (Optional) | 28 mL |  |
| TET Balance Buffer Concentrate (Optional) | 4 mL |  |
| 30X Feed Extraction Buffer (Optional) | 10 mL |  |
| 10XPBS | 25 mL |  |

If you are not planning to use the kit for over 1 month, storing Tetracycline Standard Stock, Tetracycline Antibody #1 and 100X HRP-Conjugated Antibody #2 at －20℃ or in a freezer is recommended.

### Sensitivity

| Sample Type | Detection Limit (ng/g or ppb) |
| --- | --- |
| Egg/Cheese/Sour Cream | 1.5 |
| Honey | 3 |
| Meat/Meat ProductsFish/Shrimp/Butter | 0.5 |
| Milk/Soured Milk/Yogurt/Curd | 0.5 |
| Reconstituted Milk Powder | 1.5 |
| Reconstituted Whey Powder | 1.5 |
| Feed | 8 |

### Specificity

| Analytes | Cross-Reactivity (%) |
| --- | --- |
| Tetracycline (TET) | 100.0 |
| Chlortetracycline | 124 |
| Oxytetracycline | 40 |
| Doxycycline | 25 |

### Required Materials Not Provided With the Kit

- Microtiter plate reader (450 nm)

- Vortex mixer, (e.g. Gneie Vortex mixer from VWR)

- 10, 20, 100 and 1000 L pipettes

- Multi-channel pipette: 50-300 L (Optional)

- n-hexane

### Warnings and Precautions

- The standards contain tetracycline. Handle with particular care.

- Do not use the kit past the expiration date.

- Do not intermix reagents from different kits or lots except for components with the same part No’s within their expiration dates. ANTIBODIES AND PLATES ARE KIT-AND LOT-SPECIFIC.

- Try to maintain a laboratory temperature of 20–25°C (68°–77°F). Avoid running assays under or near air vents, as this may cause excessive cooling, heating and/or evaporation. Also, do not run assays in direct sunlight, as this may cause excessive heat and evaporation. Cold bench tops should be avoided by placing several layers of paper towel or some other insulation material under the assay plates during incubation.

- Make sure you are using only distilled or deionized water since water quality is very important.

- When pipetting samples or reagents into an empty microtiter plate, place the pipette tips in the lower corner of the well, making contact with the plastic.

- Incubations of assay plates should be timed as precisely as possible. Be consistent when adding standards to the assay plate. Add your standards first and then your samples.

- Add standards to plate only in the order from low concentration to high concentration as this will minimize the risk of compromising the standard curve.

- Always refrigerate plates in sealed bags with a desiccant to maintain stability. Prevent condensation from forming on plates by allowing them equilibrate to room temperature (20 – 25℃ / 68 – 77℉) while in the packaging.

SAMPLE PREPARATION

Be sure samples are properly stored. In general, samples should be refrigerated at 2-4℃ for no more than 1-2 days. Freeze samples to a minimum of -20℃ if they need to be stored for a longer period. Frozen samples can be thawed at room temps (20 – 25℃ / 68 – 77℉) or in a refrigerator before use.

1 . Preparation of 1X OXYTET Extraction Buffer

Mix 1 volume of 5X OXYTET Extraction Buffer with 4 volumes of distilled water.

2 . Preparation of 1X TET Sample Diluent

Mix 1 volume of 10X TET Sample Diluent with 9 volumes of distilled water.

3 . Preparation of 1X TET Balance Buffer

Combine 300 μL of TET Balance Buffer Concentrate and 29.7 mL of 1X TET Sample Diluent.

4. Preparation of 1X Feed Extraction Buffer

Mix 1 volume of 30X Feed Extraction Buffer with 29 volumes of distilled water.

5 . Preparation of 1X PBS

Mix 1 volume of 10X PBS with 9 volumes of distilled water.

### Feed

- Take 1 g of feed sample, add 4 mL of 1X Feed Extraction Buffer.

- Vortex samples for 15 minutes using a multi-vortexer.

- Centrifuge for 10 minutes at 4,000 x g.

- Transfer 25 μL of clear supernatant to a new tube containing 975 μL of 1X PBS.

- Vortex samples for 1 minute.

- Use 75 μL of the sample for the assay.

Note: Dilution factor: 160

### Cheese/Egg/Sour Cream

- Add 3mL of 1X OXYTET Extraction Buffer to 1 g of sample, vortex for 10 minutes in a multi-tube vortexer or shake 30 minutes on a shaker.

- Centrifuge for 10 min at 4000 rpm.

- Transfer 200 L of supernatant to a new tube containing 300L of 1× TET Sample Balance buffer, Vortex for 1 minute.

- Use 75 L per well in the assay.

Note: Dilution factor: 10

### Honey

    - Weigh out 0.5 g of honey in a centrifuge vial.

    - Add 3.5mL 1× PBS.

    - Ensure the sample is fully dissolved in the PBS by heating for 5 – 10 minutes at 50 °C.

    - Vortex for 10 minutes in a multi-tube vortexer or shake 30 minutes on a shaker.

    - Centrifuge for 10 minutes at 4,000 x g.

    - Transfer 200 μL of the supernatant to a new tube containing 300 μL of 1X PBS.

    - Vortex for 1 minute.

    - Use 75 L of the sample for the assay.

Note: Dilution factor: 20.

### Meat/Meat ProductsFish/Shrimp/Butter

- To 1 g of homogenized sample, add 3 mL of 1X OXYTET Extraction Buffer and 1 mL n-hexane. For shrimp samples ONLY: DO NOT add n-hexane.

- Vortex for 10 minutes in a multi-tube vortexer or shake 30 minutes on a shaker.

- Centrifuge for 10 minute at 4,000 x g.

- Transfer 200 μL of the supernatant to a new tube containing 300 μL of 1X TET Sample Balance Buffer.

- Vortex for 1 minute.

- Use 75 μL per well in the assay.

Note: Dilution factor: 10

### Milk/Soured Milk/Yogurt/Curd

- To 1 g of sample (or 1 mL of liquid sample), add 3 mL of 1X OXYTET Extraction Buffer.

- Vortex for 10 minutes in a multi-tube vortexer or shake 30 minutes on a shaker.

- Centrifuge for 10 minutes at 4,000 x g.

- Transfer 200 μL of the supernatant to a new tube containing 300 L of 1X TET Sample Balance Buffer.

- Vortex for 1 minute.

- Use 75 μL per well in the assay.

Note: Dilution factor: 10

### Milk Powder

- Reconstitute 1 g of dry milk powder with 9 mL of deionized or distilled water. Mix well.

- Use this solution as the starting sample.

- To 1 mL of sample, add 3 mL of 1X OXYTET Extraction Buffer.

- Vortex for 10 minutes in a multi-tube vortexer or shake 30 minutes on a shaker.

- Centrifuge for 10 minutes at 4,000 x g.

- Transfer 200 μL of the supernatant to a new tube containing 300 μL of 1X TET Sample Balance Buffer.

- Vortex for 1 minute.

- Use 75 μL per well in the assay.

Note: Dilution factor: 10

### Whey Powder

- Reconstitute 1 g of dry whey powder with 12 mL of deionized or distilled water. Mix well.

- Use this solution as the starting sample

- To 1 mL of sample, add 3 mL of 1X OXYTET Extraction Buffer.

- Vortex for 10 minutes in a multi-tube vortexer or shake 30 minutes on a shaker.

- Centrifuge for 10 minutes at 4,000 x g.

- Transfer 200 μL of the supernatant to a new tube containing 300 μL of 1X TET Sample Balance Buffer.

- Vortex for 1 minute.

- Use 75 μL per well in the assay.

Note: Dilution factor: 10

TETRACYCLINE (TET) ELISA TEST KIT PROTOCOL

### Reagent Preparation

IMPORTANT: All reagents should be brought up to room temperature before use (1 – 2 hours at 20 – 25C / 68 – 77F); Make sure you read “Warnings and Precautions” section. Solutions should be prepared just prior to ELISA test. All reagents should be mixed by gently  inverting or swirling prior to use. Prepare volumes that are needed for the number of wells being run. Do not return the reagents to the original stock tubes/bottles. Using disposable reservoirs when handling reagents can minimize the risk of contamination and is recommended.

- Preparation of 1X Wash Solution

Mix 1 volume of the 20X Wash Solution with 19 volumes of distilled water.

- Preparation of 1X HRP-Conjugated Antibody #2

Mix 1 volume of the 100X HRP-Conjugated Antibody #2 with 99 volumes of Ab#2 Diluent.

- Special Notes for Optimal ELISA Performance

1) Allow the entire kit to equilibrate at room temperature for at least two hours before starting any ELISA assay.

2) Avoid light as much as possible during sample preparation and ELISA assay.

3) For plate washing steps: after addition of 250 μL wash buffer to the wells, incubate the plate for 20 – 30 seconds; shake the plate gently before pouring out the wash buffer. Repeat this procedure for each of the three washes.

4) Pipette all reagents and samples very accurately, especially the samples, even if you must slow down while pipetting.

### ELISA Testing Protocol

Label the individual strips that will be used and aliquot reagents as the following example:

| Component | Volume per Reaction | 24 Reactions |
| --- | --- | --- |
| Tetracycline Antibody #1 | 100 L | 2.4 mL |
| 1X HRP-Conjugated Antibody #2 | 150 L | 3.6 mL |
| 1X Wash Solution | 2.5 mL | 60 mL |
| Stop Buffer | 100 L | 2.4 mL |
| TMB Substrate | 100 L | 2.4 mL |

    - Add 75 L of each Tetracycline Standards (Negative Control, 0.05, 0.15, 0.3. 0.6, 1.5 ppb) in duplicate into different wells ( Add standards to plate only in the order from low concentration to high concentration).

    - Add 75 L of each sample in duplicate into different sample wells.

    - Add 100 L of Antibody #1 and mix well by gently rocking the plate manually for 1 minute.

    - Incubate the plate for 55 minutes at room temperature (20 – 25℃ / 68 – 77℉).

    - Wash the plate 3 times with 250 L of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels ( Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

    - Add 150 L of 1X Antibody #2 solution. Incubate the plate for 25minutes at room temperature (20 – 25C / 68 – 77F) ( Avoid direct sunlight and cold bench tops during the incubation. Covering the microtiter plate while incubating is recommended).

    - Wash the plate 3 times with 250 of L 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels ( Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

    - Add 100 L of TMB Substrate to each well. Time the reaction immediately after adding the substrate. Mix the solution by gently rocking the plate manually for 1 minute while incubating , Incubate the plate for 15 minutes at room temperature (20 – 25℃ / 68 – 77℉). ( Do not put any substrate back to the original container to avoid any potential contamination. Covering the microtiter plate while incubating is recommended).

    - After incubating,add 100 L of Stop Buffer to stop the enzyme reaction.

    - Read the plate as soon as possible following the addition of Stop Buffer on a plate reader with 450 nm wavelength ( Before reading, use a lint-free wipe on the bottom of the plate to ensure no moisture or fingerprints interfere with the readings).

### Tetracycline (TET) Concentration Calculations

A standard curve can be constructed by plotting the mean relative absorbance (%) obtained from each reference standard against its concentration in ng/mL on a logarithmic curve.

Relative absorbance (%) =

The following figure is a typical tetracycline standard curve.

REAGEN LLC

Add.: 102CommerceDr.Unit8,Moorestown,NJ08057

Tel.:856-727-0250        FAX.:856-727-0251

E-mail:reagenkits@gmail.com   reagenllc@gmail.com

Web.: www.reagen.us