REAGEN™

Neomycin ELISA Test Kit

RND99003

### Product Description

REAGEN™ Neomycin ELISA Test Kit is a competitive enzyme immunoassay for the quantitative analysis of Neomycin in meat/liver/kidney, cell lysate, urine ,serum and milk.

The unique features of the kit are:

- High recovery (>80%), rapid (10-40 minutes), and cost-effective extraction methods.

- High sensitivity (1.5 ng/g or ppb).

- High reproducibility.

- A quick ELISA assay (less than 2 hours regardless of number of samples).

### Procedure Overview

The method is based on a competitive colorimetric ELISA assay. The drug of interest has been coated in the plate wells. During the analysis, sample is added along with the enzyme-conjugate primary antibody specific for the target drug. If the target is present in the sample, it will compete for the antibody, thereby preventing the antibody from binding to the drug attached to the well. The enzyme-conjugate, will bind with the primary antibody that is complexed to the drug coated on the plate wells. The resulting color intensity, after addition of substrate, has an inverse relationship with the target concentration in the sample.

### Kit Contents, Storage and Shelf Life

REAGEN™ Neomycin ELISA Test Kit has the capacity for 96 determinations or testing of 42 samples in duplicate (assuming 12 wells for standards). Return any unused microwells to the foil bag and reseal them with the desiccant provided in the original package. Store the kit at 2-8C *. The shelf life is 12 months when the kit is properly stored.

| Kit Contents | Amount | Storage |
| --- | --- | --- |
| Neomycin-coated Microtiter Plate | 1 x 96-well Plate | 2-8C |
| Neomycin Standard Stock: Negative control 1.5 ng/mL 3 ng/mL 6ng/mL 12ng/mL 36ng/mL | 0.8 mL | 2-8C |
| Neomycin Spiking Stock: 5000 ng/mL | 0.8mL | 2-8C |
| Neomycin Antibody #1(powder) | 3 tubes |  |
| 100X HRP-Conjugated Antibody #2 | 125 uL |  |
| Antibody #2 Diluent | 20 mL |  |
| 20X Wash Solution | 28 mL |  |
| Stop Buffer | 14 mL |  |
| TMB Substrate | 12 mL |  |
| 10X Sample Extraction Buffer | 25 mL |  |
| 10X Neomycin Extraction Buffer I(Optional) | 28mL | 2-8C |
| 2X Neomycin Extraction Buffer II(Optional) | 28mL | 2-8C |
| 1X Neomycin Milk Clean Up Reagent (Optional) | 10mL | 2-8C |
| 10X Neomycin Balance Buffer (Optional) | 10ml | 2-8C |

*   If you are not planning to use the kit for over 1 month, store Neomycin Standard Stock, Neomycin Antibody #1and 100X HRP-Conjugated Antibody #2 at -20C or in a freezer.

### Sensitivity

| Sample Type | Detection Limit (ng/g or ppb) |
| --- | --- |
| Fish/Shrimp/Meat/liver/kidney | 30 |
| Cell Lysate | 15 |
| Urine /Serum | 40 |
| Milk | 30 |
| Honey | 60 |
| Feed | 30 |

### Specificity

| Analytes | Cross-Reactivity (%) |
| --- | --- |
| Neomycin | 100 |
| Gentamicin | <0.1 |
| Kanomycin | <0.1 |
| Streptomycin | <0.1 |
| Tobramycin | <0.1 |
| Amikazin | <0.1 |
| Dihydrostreptomy | <0.1 |

### Required Materials Not Provided With the Kit

- Microtiter plate reader (450 nm)

- Tissue Mixer (e.g. Omni TissueMaster Homogenizer)

- Vortex mixer (e.g. Genie Vortex mixer from VWR)

- 10, 20, 100 and 1000 uL pipettes

- Multi-channel pipette: 50-300 uL (Optional)

- n-Hexane

### Warnings and Precautions

- The standards contain Neomycin. Handle with particular care.

- Do not use the kit past the expiration date.

- Do not intermix reagents from different kits or lots except for components with the same part No’s within their expiration dates. ANTIBODIES AND PLATES ARE KIT-AND LOT-SPECIFIC. Make sure that the antibody #2 and diluent are mixed in correct volumes.

- Try to maintain a laboratory temperature of 20°–25°C (68°–77°F). Avoid running assays under or near air vents, as this may cause excessive cooling, heating and/or evaporation. Also, do not run assays in direct sunlight, as this may cause excessive heat and evaporation. Cold bench tops should be avoided by placing several layers of paper towel or some other insulation material under the assay plates during incubation.

- Make sure you are using only distilled or deionized water since water quality is very important

- When pipetting samples or reagents into an empty microtiter plate, place the pipette tips in the lower corner of the well, making contact with the plastic.

- Incubations of assay plates should be timed as precisely as possible. Be consistent when adding standards to the assay plate. Add your standards first and then your samples.

- Add standards to plate only in the order from low concentration to high concentration as this will minimize the risk of compromising the standard curve.

- Always refrigerate plates in sealed bags with a desiccant to maintain stability. Prevent condensation from forming on plates by allowing them equilibrate to room temperature (20 – 25C / 68 – 77F) while in the packaging.

SAMPLE PREPARATION

Be sure samples are properly stored. In general, samples should be refrigerated at 2-4°C for no more than 1-2 days. Freeze samples to a minimum of -20°C if they need to be stored for a longer period. Frozen samples can be thawed at room temps (20 – 25C / 68 – 77F) or in a refrigerator before use.

1. Preparation of 1X Sample Extraction Buffer:

Mix 1 volume of 10X Sample Extraction Buffer with 9 volumes of distilled water.

2. Preparation of 1X Neomycin Extraction Buffer I:

Mix 1 volume of 10X Neomycin Extraction Buffer I with 9 volumes of distilled water.

3. Preparation of 1X Neomycin Extraction Buffer II:

Mix 1 volime of 2X Neomtcin Extraction Buffer II with 1 volume of ditilled water.

4.Preparation of 1X Neomycin Balance Buffer：

Mix 1 volume of 10X Neomycin Balance Buffer with 9 volume of distilled water

### Cell Lysate

- Take 100 uL of cell lysate, add 900 uLof 1X Sample Extraction Buffer, mix well.

- Centrifuge for 5 minutes at 4000 x g.

- Take 50 uL of the supernatant per well for the assay.

Note: Dilution factor: 10.

### Shrimp/Fish/ Meat/Liver/Kidney

- To 1.0 g of the homogenized sample, add 3.5mL of 1X Neomycin Extraction Buffer I, 500μL of 1X Neomycin Extraction Buffer II and 3 mL of hexane.

- Vortex the sample for 3 minutes manually or vortex the sample with multi-vortexer for 10 minutes..

- Centrifuge the sample for 10 minutes at 4,000 x g.

- Remove the top hexane layer completely. Carefully transfer 250μL of the aqueous layer to a new tube (avoid any contamination from the hexane layer!), add 250μL of 1X Sample Extraction Buffer and vortex for 30 seconds.

- Use 50 μL per well for the assay.

Note:Dilution factor: 10.

### Urine/Serum

- Take 50uL of urine or serum sample, add 1.2 mL of 1X Sample Extraction Buffer, mix well.

- Centrifuge for 5 minutes at 4,000 x g.

- Take 50uL of the supernatant per well for the assay.

Note:Dilution factor: 25.

### Milk

- Add 1 mL of milk sample into a centrifuge tube.

- Add 125 uL of 1X Neomycin Milk Clean Up Reagent.

- Vortex vigorously for 20 seconds.

- Centrifuge at 4,000 x g for 10 minutes.

- Take 56 uLof supernatant and 944 uL 1X Sample Extraction Buffer, mix well.

- Use 50uL of the sample for the assay.

Note:Dilution Factor:20

### Feed

- To 1.0 g of the homogenized sample, add 7.0 mL of 1X Neomycin Extraction Buffer I, 700 uL of 1X Neomycin Extraction Buffer II and 3 mL of hexane.

- Vortex the sample for 3 minutes manually or vortex the sample with multi-vortexer for 10 minutes.

- Centrifuge the sample for 10 minutes at 2,000 x g.

- Remove the top hexane layer completely. Carefully transfer 200 uLof the aqueous layer to a new tube (avoid any contamination from the hexane layer!), add 800 uL of 1X Neomycin Balance Buffer and vortex for 30 seconds.

- Use 50 uL per well for the assay.

Note: Dilution factor: 40.

Honey

1. To 1.0 g of the homogenized sample, add 4.2 mL of 1X Neomycin Extraction Buffer I,

800 uL of 1X Neomycin Extraction Buffer II.

2. Heat at 70–80℃for 5 minutes or until the honey will dissolve in solution.

3. Vortex the sample for 1 minutes manually.

4. Centrifuge the sample for 10 minutes at 4,000 x g.

5. Carefully transfer 300 uLsupernatant to 700 uL of 1X Neomycin Balance Buffer and

vortex for 30 seconds.

6. Use 50 uL per well for the assay.

Note: Dilution factor: 20.

NEOMYCIN ELISA TEST KIT PROTOCOL

### Reagent Preparation

IMPORTANT: All reagents should be brought up to room temperature before use (1 – 2 hours at 20 – 25C / 68 – 77F); Make sure you read “Warnings and Precautions” section on page 3. Solutions should be prepared just prior to ELISA test. All reagents should be mixed by gently inverting or swirling prior to use. Prepare volumes that are needed for the number of wells being run. Do not return the reagents to the original stock tubes/bottles. Using disposable reservoirs when handling reagents can minimize the risk of contamination and is recommended.

- Preparation of 1X HRP-Conjugated Antibody #2

Mix 1 volume of 100X HRP-Conjugated Antibody #2 with 99 volumes of Antibody #2  Diluent

- Preparation of 1X Wash Solution

Mix 1 volume of the 20X Wash Solution with 19 volumes of distilled water

- Preparation of Neomycin Antibody#1 Solution

Based on the needed amount of antibody, take out one or more Neomycin Antibody#1 tubes   (powder), add 2 mL of 1X Sample Extraction Buffer to each tube needed, vortex 30 seconds, leave the tube at room temperature for at least 15 min, vortex 30 seconds. The antibody solution is ready to use. The reconstituted antibody is stable for 1 week at 4C or 1 month at -20C freezer.

### ELISA Testing Protocol

Label the individual strips that will be used and aliquot reagents as the following example:

| Component | Volume per Reaction | 24 Reactions |
| --- | --- | --- |
| Neomycin Antibody #1 | 50uL | 1.2mL |
| 1X HRP-Conjugated Antibody #2 | 50uL | 1.2mL |
| 1X Wash Solution | 2mL | 48mL |
| Stop Buffer | 100 uL | 2.4 mL |
| TMB Substrate | 100 uL | 2.4 mL |

- Add 50 uL of each Neomycin Standards in duplicate into different wells (Add standards to plate only in the order from low concentration to high concentration).

- Add 50 uL of each sample in duplicate into different sample wells.

- Add 50 uL of Antibody#2 to each well and 50 uL of Antibody #1 to each well. Mix well by gently rocking the plate manually for 1 minute.

- Incubate the plate for 30 minutes at room temperature (20 – 25C / 68 – 77F). (Avoid direct sunlight and cold bench tops during the incubation. Covering the microtiter plate while incubating is recommended).

- Wash the plate 3 times with 250 uL of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels (Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

- Add 100 uLof TMB substrate.Time the reaction immediately after adding the substrate.Mix the solution by gently rocking the plate manually for 1 minute while incubating(Do not put any substrate back to the original container to avoid any potential contamination. Any substrate solution exhibiting coloration is indicative of deterioration and should be discarded. Covering the microtiter plate while incubating is recommended).

- After incubating for 15 minutes at room temperature (20 – 25C / 68 – 77F), add 100 uL of Stop Buffer to stop the enzyme reaction.

- Read the plate as soon as possible following the addition of Stop Buffer on a plate reader with 450 nm wavelength (Before reading, use a lint-free wipe on the bottom of the plate to ensure no moisture or fingerprints interfere with the readings).

### Neomycin Concentration Calculations

A standard curve can be constructed by plotting the mean relative absorbance (%) obtained from each reference standard against its concentration in ng/mL on a logarithmic curve.

Relative absorbance (%) =

Use the mean relative absorbance values for each sample to determine the corresponding concentration of the tested drug in ng/mL from the standard curve.

The following figure is a typical neomycin standard curve.

REAGEN LLC

Add.: 102CommerceDr.Unit8,Moorestown,NJ08057

Tel.:856-727-0250       FAX.:856-727-0251

E-mail:reagenkits@gmail.com  reagenllc@gmail.com

Web.:www.reagen.us