REAGEN™

Glucose Assay Kit

RNF91001

Product Description

REAGEN™ Glucose Assay Kit is a cuvette-based enzymatic assay for the determination of Glucose across a variety of food samples. The kit uses a spectrophotometric, kinetic assay to detect the enzymatic degradation of Glucose directly from food samples.

The unique features of the kit are:

- High sensitivity and low detection limit (0.8 ppm)

- High reproducibility

- A rapid (30 minute), robust enzyme-based assay

Procedure Overview

REAGEN™ Glucose Assay Kit uses a coupled enzymatic reaction to detect glucose and/or fructose in food samples: glucose is detected by phosphorylation with hexokinase to make glucose-6-phosphate (G-6-P). This compound is then used as a substrate by the enzyme glucose-6-phosphate dehydrogenase to convert NADP+ to NADPH. This conversion, measured at 340 nm, is proportional to the level of glucose in the sample.

The absorbance of each sample at 340 nm is measured using a spectrophotometer. The concentration of glucose in each sample is then directly determined from the change in absorbance at 340 nm within 30 minutes.

Kit Contents, Storage and Shelf Life

REAGEN™ Glucose Assay Kit has the capacity for 48 determinations or 24 determinations in duplicate. Store the kit at 4°C. The shelf life of the kit is 12 months, after receipt, when the kit is properly stored.

| Kit Contents | Amount | Storage |
| --- | --- | --- |
| Dilution Buffer | 10 mL | 4ºC |
| Substrate Mix | 20 mL | 4ºC |
| Hexokinase | 0.6 mL | 4ºC |
| Microtiter Plate | One | 4ºC |

Sensitivity (Detection Limit)

| Sample Volume | Detection Limit (mg/L) |
| --- | --- |
| 0.01 mL | 8 |
| 0.1mL | 0.8 |

The linear range of the assay is 0.8 – 1000 mg/L (ppm).

Required Materials Not Provided With the Kit

- Spectrophotometer (340 nm)

- Low volume (0.3 mL) UV cuvettes, such as Brand Scientific catalog #7592 00

- Deionized or distilled water

- 50 mL plastic conical tubes

Warnings and Precautions

- Do not use the kit past the expiration date.

- Try to maintain a laboratory temperature of 20°–25°C (68°–77°F). Avoid running assays under or near air vents, as this may cause excessive cooling, heating and/or evaporation. Also, do not run assays  in  direct sunlight,  as this may cause excessive  heat and evaporation. Cold  bench tops should be avoided by placing several layers of paper towel or some other insulate material under the assay plates during incubation.

- Make sure you are using only distilled or deionized water since water quality is very important.

- When pipetting samples or reagents  into an empty microtiter plate, place the pipette tips in the lower corner of the well, making contact with the plastic.

SAMPLE PREPARATION

General information on sample preparation

- Grind or homogenize solid samples then extract with water and filter to clarify samplexer.

- Adjust acidic samples to pH 7.5 using NaOH solution. Allow pH of carbonated samples to stabilize for 15 minutes before final pH adjustment.

- Measure colored samples against a sample blank containing only buffer plus sample (using the same sample volume used for the assay). Strongly colored samples can be treated with polyvinylpolypyrrolodine (PVPP, final concentration about 1%) to decolorize: Stir 1 minute and filter to remove solids.

- Extract samples containing fat with hot water. Cool the extract and remove the aqueous sample from the fat layer.

- Protein rich samples such as milk-based samples should be deproteinized before performing the assay. Use a Carrez based deproteinization since, unlike acid-based deproteinization procedures, it will not hydrolyze polysaccharides in the sample

Carrez deproteinization: Transfer 0.5 – 5 g sample to a 50 mL volumetric flask containing about 30    mL dH2O. Add 2.5 mL Carrez I solution (3.6% potassium hexacyanoferrate (II) • 3 H2O) and 2.5 mL Carrez II solution (7.2% zinc sulfate • 7 H2O). Mix and carefully adjust the pH to 8 with sodium hydroxide, adjust the volume to 50 mL, mix and filter.

### Beer

To remove the carbonation (bubbles), stir 10 mL of beer in a beaker for 30 – 40 seconds with a glass rod or filter using a fluted filter paper. The “decarbonated” sample can be used undiluted in the assay

### Fruit Juices

Juice samples can be treated with polyvinylpolypyrrolodine (PVPP, final concentration about 5%) to decolorize: Stir 1 minute and filter to remove solid

### Fruit Jam

Accurately weigh about 0.5 g sample and transfer to a 100 mL volumetric flask containing about 99 mL dH2O. Mix, adjust to 100 mL volume and filter

### Condensed Milk

Accurately weigh about 0.5 g sample and transfer to a 50 mL volumetric flask   containing about 30 mL dH2O. Mix and incubate 10 minutes at room temperature. Add 2.5 mL Carrez I solution (3.6% potassium hexacyanoferrate (II) • 3 H2O) and 2.5 mL Carrez II solution (7.2% zinc sulfate • 7 H2O). Mix and carefully adjust the pH to 8 with sodium hydroxide, adjust the volume to 50 mL, mix and filter to clarif

### Ice Cream

Accurately weigh about 0.5 g sample and transfer to a 100 mL volumetric flask c ontaining about 99 mL dH2O. Mix, adjust to 100 mL volume and filte

### Potato-based Foods, Fruits and Vegetables

Homogenize > 25 g sample (using a blender or other homogenizer). Transfer 10 g to  a 50 mL tube and add 30 mL dH2O. Vortex mix or shake for 3 minutes. Add 1.25 mL   Carrez I solution (3.6% potassium hexacyanoferrate (II) • 3 H2O) and 1.25 mL   Carrez II solution (7.2% zinc sulfate • 7 H2O). Mix and carefully adjust the pH to 7.5 with sodium hydroxide, add 50 µL n-butanol and adjust the volume to volume to 50 mL

### Honey

Dilute the honey 1:1000 in dH2O and use in the assay.

GLUCOSE DETECTION PROTOCOL

### Reagent Preparation

IMPORTANT: All reagents should be brought up to room temperature before use (1 – 2 hours at20 – 25°C / 68 – 77°F); Make sure you read “Warnings and Precautions” section on page 2.

### ELISA Testing Protocol

Note: Each determination is performed in 0.3 mL microcuvettes, in parallel, using      samples diluted into water without Glucosese enzyme. The color change of these no-enzyme controls (Sample blank) is then subtracted from the color change of corresponding enzyme reactions (Sample reaction) to correct for sample background signals.

| Determination of Glucose in Food Samples Pipet into low-volume cuvettes | Sample blank | Sample Reaction |
| --- | --- | --- |
| Sample Dilution Buffer Note: Sample volume + Buffer volume should equal 100 µL. If the sample is not strongly colored, the sample volume may be >10 µL. Sample + Buffer = 100 µL. Substrate Mix | ---- 100 µL 190 µL | 10 µL 90 µL 190 µL |
| Mix and immediately measure the absorbance at 340 nm (A1). |  |  |
| Hexokinase | 10 µL | 10 µL |
| Incubate 15 min at 37°C or 30 min at room temperature. Measure the absorbance at 340 nm (A2). |  |  |

To calculate the glucose concentration, use the following general equation:

c =  V x MW x ΔA  (g / L)

ε x d x v

V = final volume (0.31 mL)

v = sample volume (0.01 – 0.1 mL)

MW = molecular weight of glucose (180.16 g/mol)

d = light path (1 cm)

ε = extinction coefficient of NADH (6220 at 340 nm)

ΔA = (A2 – A1)sample glucose – (A2 – A1)sample blank

For example for a 10 µL sample:

c =  0.31 x 180.16     x ΔA (g / L) = 0.898 x ΔA (g glucose / L sample solution)

6220 x 1 x 0.01

# Alterative Protocol using a 96-well microplate

Note: The microplate protocol uses a reaction volume of 300 μL.

- Add 10 μL to 100 μL of sample solution, depending on the sample type, in triplicate into microplate wells.

- Add Dilution Buffer into the wells for a total volume of 100 μL in each well.

- Add 190 μL of Substrate Mix into each well.

- Immediately measure the absorbance of each sample at 340 nm (A1).

- Add 10 μL of Hexokinase into two of the three of triplicate wells; add 10 μL of dH2O into the other control (Sample Blank) well. Mix

- Incubate at room temperature; after 30 min read the absorbance at 340nm (A2).

Data analysis is performed in the manner described in the previous section.

REAGEN LLC

Add.: 102CommerceDr.Unit8,Moorestown,NJ08057

Tel.:856-727-0250         FAX.:856-727-0251

E-mail:reagenkits@gmail.com  reagenllc@gmail.com

Web.:www.reagen.us