REAGEN™

SALMONELLA ELISA Test Kit

RNB93001

### Product Description

REAGEN™Salmonella ELISA Test Kit is designed for routine screening for Salmonella from cultured food and feed samples. Salmonella is a genus of gram negative bacteria belonging to the Enterobacteriaceae family. As a major contributor to foodborne illness, Salmonella is founded in poultry, meat and eggs. When Salmonella is consumed it may result in the disease known as Salmonellosis. The detection of Salmonella in food samples is often laborious and time-consuming. Several days are required to culture the bacteria from a sample followed by a final isolation and identification by biochemical and serological techniques. The REAGEN™Salmonella ELISA Test Kit provides a simple, rapid, sensitive and cost-effective enzyme immunoassay (ELISA) screening method, which enables government agencies,seafood processors,as well as quality assurance organizations,to detect positive and negative Salmonella samples to as low as 105 cells/mL level.

### Procedure Overview

The method is based on a colorimetric ELISA assay. The monoclonal antibody against Salmonella has been coated in the plate wells. During the analysis, the sample is added. If Salmonella is present in the sample, it will bind with the antibody.A suitable broth is added to the well for rapid growth.The secondary antibody, tagged with a peroxidase enzyme, targets the Salmonella that is complexed to the primary antibody coated on the plate wells. The resulting color intensity, after addition of substrate, is related to the amount of Salmonella in the sample.

### Kit Contents, Storage and Shelf Life

REAGEN™Salmonella ELISA Test Kit has the capacity for 96 determinations or testing of 46 samples in duplicate (assuming 4 wells for positive and negative controls). Most of the kit components should be stored at 2-8 C except for the Salmonella controls and 50X Ab#2 stock which should be stored at -20 C. Return any unused microwells to the foil bag and reseal them with the desiccant provided in the original package. The shelf life is 12 months when the kit is properly stored.

| Kit Contents | Amount | Storage |
| --- | --- | --- |
| Salmonella antibody-coated Plate | 1x 96-well Plate(8 wells x 12 strips) | 2-8C |
| 1X Salmonella Negative Control | 1.0 mL | -20C |
| 1X Salmonella Positive Control | 1.0 mL | -20C * |
| HRP-Conjugated Antibody #2 (powder) | 3 tubes | -20C |
| Antibody #2 Diluent 20X Wash Solution Stop Buffer TMB Substrate | 20 mL 30 mL 20 mL 10 mL | 2-8C |

### Sensitivity

The kit can detect 105cells/mL

### Specificity

The kit can broadly detect Salmonella serotypes of group A, B, C, D, E and other Salmonella antigen classes. The intensity of reaction may vary with the strain of Salmonella tested

### Required Materials Not Provided With the Kit

- Microtiter plate reader (450 nm)

- Incubator

- Vortex mixer (e.g. Gneie Vortex mixer from VWR)

- 10, 20, 100 and 1000 L pipettes

- Multi-channel pipette: 50-300 L (Optional)

- Modified GN Broth: GN broth containing 10µg/ml Novobiocin. GN broth can be purchased or prepare in-house per the following protocol.

- Gram Negative Broth (GN Broth or Hajna GN Broth, pH 7.0 ± 0.2)：

Tryptose . . . . . . . . . . . .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .. . . . . . . . . . . . . . . . .. . . . . 20.0g

Dex. . . . . . . . . . . .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 1.0g D-Mannitol . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .. . . . . . . . . . . . . . . . . . . . . . . . 2.0g  Sodiumcitrate . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 5.0 g Sodium desoxycholate . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 0.5 g Dipotassium phosphate . . . . . . . . . . . . . . . . . . . . . . . .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . 4.0 g Monopotassium phosphate . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 1.5 g Sodium chloride . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 5.0 g Distilled water . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  1.0 L

（Dissolve ingredients in water. Autoclave at 121°C for 15 min. Avoid overheating）

- Autoclave or boiling water bath (100°C)

- Culture bottles (500ml)

- Microbiological media for enrichment of Salmonella

- Microbiological media for the confirmation of Salmonella

- Distilled water

### Warnings and Precautions

- The samples contain Salmonella. Handle with particular care.

- Good microbiological practices (i.e. aseptic technique, safety precautions) should be employed when carrying out sample preparation and enrichment procedures.

- Do not use the kit past the expiration date.

- Do not intermix reagents from different kits or lots except for components with the same part No’s within their expiration dates.

- Try to maintain a laboratory temperature of 20°–25°C (68°–77°F). Avoid running assays under or near air vents, as this may cause excessive cooling, heating and/or evaporation. Also, do not run assays in direct sunlight, as this may cause excessive heat and evaporation. Cold bench tops should be avoided by placing several layers of paper towel or some other insulation material under the assay plates during incubation.

- Make sure you are using only distilled or deionized water since water quality is very important.

- When pipetting samples or reagents into an empty microtiter plate, place the pipette tips in the lower corner of the well, making contact with the plastic.

- Incubations of assay plates should be timed as precisely as possible. Be consistent when adding standards to the assay plate. Add your standards first and then your samples.

- Add standards to plate only in the order from low concentration to high concentration as this will minimize the risk of compromising the standard curve.

- Always refrigerate plates in sealed bags with a desiccant to maintain stability. Prevent condensation from forming on plates by allowing them equilibrate to room temperature (20 – 25C / 68 – 77F) while in the packaging.

SAMPLE PREPARATION

Be sure samples are properly stored. In general, samples should be refrigerated at 2-4°C for no more than 1-2 days. Freeze samples to a minimum of -20°C if they need to be stored for a longer period. Frozen samples can be thawed at room temps (20 – 25C / 68 – 77F) or in a refrigerator before use.

### Food sample

Add 25 g of homogeneous food sample or 25 mL liquid sample to 225 mL of pre-enrichment broth (e.g. buffered peptone water, BPW). The sample may require blending. Incubate for 18-24 hours at 37°C. The broth may be stored refrigerated for up to 48 hours.

SALMONELLA ELISA TEST KIT PROTOCOL

### Reagent Preparation

IMPORTANT: All reagents should be brought up to room temperature before use (1 – 2 hours at 20 – 25C / 68 – 77F); Make sure you read “Warnings and Precautions” section on page 3. Solutions should be prepared just prior to ELISA test.  All reagents should be mixed by gently inverting or swirling prior to use. Prepare volumes that are needed for the number of wells being run. Do not return the reagents to the original stock tubes/bottles. Using disposable reservoirs when handling reagents can minimize the risk of contamination and is recommended.

- Preparation of 1X Wash Solution

Mix 1 volume of 20X Wash Buffer concentrate with 19 volumes of distilled water.

- Preparation of 1X HRP-Conjugated Antibody #2

Based on the needed amount, take out one or more HRP-Conjugated Antibody #2 tubes (powder), add 2.5 mL of Antibody #2 Diluent to each tube, invert the tube up and down for 10 times, leave the tube at room temperature for at least 15 min, invert the tube up and down for another 10 times. The antibody#2 solution is ready to use. The reconstituted antibody#2 is stable for 1 week at -20 C freezer.

### ELISA Testing Protocol

Label the individual strips that will be used and aliquot reagents as the following example:

| Component | Volume per Reaction |
| --- | --- |
| 1X Negative Control | 50L |
| 1X Positive Control | 50 L |
| 1X HRP-Conjugated Antibody #2 | 50 L |
| 1X Wash Solution | 2.0 mL |
| Stop Buffer | 100 L |
| TMB Substrate | 100 L |

    - Take out the needed plate strips, add 50 L of Negative and Positive Control in duplicate into different wells

    - Add 50L of each sample in duplicate into different sample wells.

    - Incubate the plate for 30 minutes at 37 C.

    - Wash the plate 3 times with 300L of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels ( Perform next step immediately after plate washings. Do not allow the plate to air dry between working steps).

Add 250 μL Modified GN Broth to each well. Cover the wells and incubate for 4 hours at 37 °C.Transfer the well broth to another container if needed for future validation (optional). Wash the plate 3 times with 300 L of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels.

    - Add 50L of 1X HRP-Conjugated Antibody #2 to each well. Incubate the plate for 30 minutes. (Avoid direct sunlight and cold bench tops during the incubation. Covering the microtiter plate while incubating is recommended).

Wash the plate 3 times with 300 L of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels (Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

    - Add 100L of TMB substrate. Time the reaction immediately after adding the substrate. Mix the solution by gently rocking the plate manually for 1 minute while incubating. ( Do not put any substrate back to the original container to avoid any potential contamination. Any substrate solution exhibiting coloration is indicative of deterioration and should be discarded. Covering the microtiter plate while incubating is recommended).

    - After incubation for 15 minutes at room temperature (20 – 25 C / 68 – 77 F), add 100 L of Stop Buffer to stop the enzyme reaction.

    - Read the plate as soon as possible following the addition of Stop Buffer on a plate reader with 450 nm wavelength (Before reading, use a lint-free wipe on the bottom of the plate to ensure no moisture or fingerprints interfere with the readings).

### Interpretation of the Results

Visual Interpretation of the Results：

Visual assessment must be done prior to addition of stopping solution.

Positive Control:	Blue

Negative Control:	Clear or very pale blue

A sample is considered positive when the following criteria are met:

1.	The negative control is clear or very pale blue in color and the positive sample is significantly deeper in color than this.

2.	The positive sample should always be significantly deeper in color than the negative control, however it need not be as blue as the positive control.

A sample is considered negative when the following criteria are met:

1.  The negative control is clear or very pale blue in color.

2.  The sample is less than or equal to the color of the negative control.

Interpret results and record as positive or negative on the record sheet provided. Return reagents to 2-8°C. Negative/Positive samples can only be interpreted if the expected results are produced with the controls.

OD Method

Measure the plate absorbance using a microplate reader at A450nm and record absorbance values on the record sheet provided.

Negative Control:Readings at A450nm should give an OD of less than 0.4.

If not this indicates insufficient washing of the microplate

Positive Control: Readings at A450nm should give an OD of greater than 1.0.

If controls are not within the limits noted above, the test must be repeated.

Negative samples:Samples that give A450nm <0.6

Positive samples:Samples that give A450nm ≥ 0.6

Negative/positive samples can only be interpreted if the expected results are produced with the controls.

If confirmation of Salmonella ELISA Test positive samples is required, we would recommend one of the following confirmation procedures.

1.	Transfer 1ml of each of the suspect pre-enriched sample broths to a fresh 10 mL aliquot of mGN

broth and incubate at 42 C for 18-24 hours. At the same time, streak each of the mGN broths to

at least two salmonella selective agar media and incubate for 18-24 hours at 35 C as instructed

by the manufacturer. Normally Xylose Lysine Desoxycholate agar (XLD), Hektoen-Enteric agar and Bismuth Sulphite agar are used.

- Examine  selective  agar  media  for  suspicious  colonies  and,  if  present,  confirm  these  using biochemical and serological testing.

REAGEN LLC

Add.: 102CommerceDr.Unit8,Moorestown,NJ08057

Tel.:856-727-0250         FAX.:856-727-0251

E-mail:reagenkits@gmail.com  reagenllc@gmail.com

Web.:www.reagen.us